2001

Quality Assurance Report

April 2003

Additional copies of this report may be obtained from:

Colorado Department of Public Health and Environment

Air Pollution Control Division

Technical Services Program

APCD-TS-B1

4300 Cherry Creek Drive South

Denver, CO 80246 -1530

(303) 692-3100

TABLE OF CONTENTS

Section											Page

LIST OF TABLES	ii

LIST OF FIGURES	iii

EXECUTIVE SUMMARY	iv

INTRODUCTION	1

1.0 STATUS OF THE MONITORING NETWORK	4

2.0 GASEOUS MONITORING NETWORK QUALITY ASSURANCE	10

2.1 Assessment and Accuracy Audits	10

2.2 Data Validation	14

2.3 EPA Interlaboratory Studies	14

2.4 Precision and Accuracy Program	17

2.5 System Audits	29

2.6 Additional Audits	29

2.7 Certification of Audit Standards	29

3.0 PARTICULATE MONITORING NETWORK QUALITY ASSURANCE	31

3.1 Accuracy Audits	31

3.2 Data Validation	36

3.3 EPA Interlaboratory Studies	36

3.4 Precision and Accuracy Program	42

3.5 Special Study Audits	48

3.6 Additional Audits	48

3.7 Certification of Audit Standards	48

4.0 AIR POLLUTION LABORATORY QUALITY ASSURANCE	49

4.1 Duplicate Samples	50

4.2 Spiked Samples	50

4.3 EPA Interlaboratory Audits	51

5.0 METEOROLOGICAL MONITORING QUALITY ASSURANCE	54

6.0 CORRECTIVE ACTION SUMMARY	55

LIST OF TABLES

Table 1 -  EPA Minimum Requirements for Data Quality	2

Table 2 -  APCD Data Quality Objectives	3

Table 3 -  2001 Gaseous Data Recovery	6

Table 4A - 2001 Particulate and Lead Data Recovery	7

Table 4B - Data Recovery for Continuous Patuculate Samplers	9

Table 5 -  2001 Gaseous Internal Audits Summary	12

Table 6 -  NPAP Interlaboratory Gaseous Audit Results	16

Table 7 -  Carbon Monoxide Precision and Accuracy Probability Limits	19

Table 8 -  Ozone Precision and Accuracy Probability Limits	21

Table 9 -  Sulfur Dioxide Precision and Accuracy Probability Limits	23

Table 10 -Nitrogen Dioxide Precision and Accuracy Probability Limits	24

Table 11 -2001 Particulate Internal Audits Summary	32

Table 12 -Bias on Accuracy Audits for PM2.5	34

Table 13 -NPAP Interlab TSP and PM10 Audit Results	37

Table 14 -PM 2.5 Interlab Performance Evaluation Program Audit Results	41

Table 15 -2001 PM10 and TSP Precision and Accuracy Probability Limits	43

Table 16 -PM2.5 Precision Data Confidence Limits	44

Table 17 -Laboratory Quality Control Checks	49

Table 18 -Laboratory Spike Recoveries for Lead	52

Table 19 -Laboratory Duplicate Analyses and NPAP Results for Lead	53

Table 20 -Laboratory Quality Control for Sulfate	53

LIST OF FIGURES

Figure 1 - Number of Gaseous APCD Accuracy Audits (1990-2001)	13

Figure 2 - APCD Annual Comparison of Carbon Monoxide Precision Probability Limits (1990 - 2001)	25

Figure 3 - Figure 3 - APCD Annual Comparison of Ozone Precision Probability Limits (1990 - 2001)	26

Figure 4 - Figure 4 - APCD Annual Comparison of Sulfur Dioxide Precision Probability Limits (1990-2001)	27

Figure 5 - Figure 5 - APCD Annual Comparison of Nitrogen Dioxide Precision Probability Limits (1991 - 2001)	28

Figure 6 - Number of APCD Particulate Accuracy Audits Performed Annually (1990 - 2001)	35

Figure 7 - APCD Annual Comparison of PM10 Precision Probability Limits (1990-2001)	45

Figure 8 - APCD Annual Comparison of TSP Precision Probability Limits (1990-2001)	46

Figure 9 - APCD PM2.5 Precision Confidence Intervals (1999-2001)	47

APPENDIX A - 2001 Site List	58

EXECUTIVE SUMMARY

Gaseous Monitoring Network

The APCD routinely monitors four gaseous pollutants: carbon monoxide (CO), ozone (O3), sulfur dioxide (SO2), and oxides of nitrogen (NOx).  Fifty-eight accuracy audits were conducted in 2001.  All but three of these audits were within the 10% full-scale analyzer response error considered acceptable by the APCD.

EPA interlaboratory comparisons were conducted during 2001 for carbon monoxide, ozone, sulfur dioxide, and nitrogen dioxide.  The ozone comparisons were invalidated due to faulty NPAP equipment.  Interlab comparisons are included for 18 gaseous analyzers.  All interlab results were within the acceptable EPA performance criterion of 15% error, except one.  It was determined that the instrument in question was in need of repair and recalibration.

Overall data recovery for the gaseous monitoring network was 92% in 2001.  During this time only one of the analyzers in the APCD network failed to meet the EPA 75% data completeness criterion.

Particulate Monitoring Network

The APCD routinely monitors three size classes of particulate matter: total suspended particulate (TSP), inhalable particulate matter (PM10), and respirable particulate matter (PM2.5).  Ambient lead concentrations are determined from analysis of TSP filters collected at selected sites.  Field audits were performed at least semi-annually at all particulate monitoring sites in Colorado.  A total of 606 particulate audits were conducted in 2001.  Ninety-four percent of these audits passed both accuracy and design flow criteria.  Only two audits during 2001 failed  to meet the EPA accuracy criterium.

Ninety-two PM10 samplers and one TSP sampler were audited under the EPA NPAP interlaboratory program in 2001.  All but 2 of these interlaboratory audits showed flow errors within EPA’s acceptable criteria.  Four PM2.5 sites were audited under the EPA FRM PEP audit program to determine network bias.  The bias of the PM2.5 network was determined to be (-)2.5% in 2001.

Overall APCD particulate network data recovery in 2001 was 93% for TSP, 85% for lead, 88% for PM10, and 93% for PM2.5.

Meteorological Monitoring Network

Audits of the meteorological monitoring network were conducted at 13 of the 19 meteorological sites during 2001.

INTRODUCTION

The Colorado Department of Public Health & Environment (CDPHE) Air Pollution Control Division (APCD) operates a statewide system of air pollution monitors to characterize air quality in Colorado.  Colorado is required to monitor air pollutants for which National Ambient Air Quality Standards (NAAQS) have been established.  These NAAQS air pollutants include four gaseous pollutants; carbon monoxide (CO), ozone (O3), sulfur dioxide (SO2), and nitrogen dioxide (NO2), and three particulate pollutants; respirable particulate (PM2.5), inhalable particulate (PM10), and lead (Pb).  In addition, a total suspended particulate (TSP) sampling network and several meteorological monitoring stations are also operated by the APCD in a number of Colorado communities.  The gaseous pollutant networks, particulate pollutant networks, and meteorological monitoring program are maintained and operated by the APCD Technical Services Program.

The data collected by this network provide valuable information used in the protection of public health and welfare.  Analysis of ambient data collected by this air monitoring network serves as the basis for determination of air pollutant trends, aids in identification of NAAQS pollutant attainment and nonattainment areas, aids in review of the air quality impacts of new and existing sources, and is used in the validation of air quality models used in the development of pollution control strategies.

The Air Pollution Control Division is required to develop and implement a quality assurance program consisting of policies, procedures, specifications, standards, and documentation necessary to; 1) provide data of adequate quality to meet monitoring objectives, and 2) minimize loss of air quality data due to sampler malfunctions or out-of-control situations.  Minimum standards for air quality monitoring programs are detailed in the Code of Federal Regulations (Part 58 Ambient Air Quality Surveillance).  More detailed information about APCD's network is provided in the APCD Annual Network Review.  More specific information regarding quality assurance and monitoring operations is available in the Quality Assurance Project Plan and the associated Standard Operating Procedures found in its appendices.  The Technical Services Program has a Quality Assurance group charged with performing or coordinating quality control efforts to meet the objectives of this monitoring effort.

The 2001 Quality Assurance Report is part of an annual assessment of quality assurance activities conducted by the APCD and is used to inform federal regulators, APCD management and staff, and the general public of those activities.  Minimum EPA Requirements for Data Quality and APCD’s Data Quality Goals are summarized in the following tables:

Table 1 - EPA Minimum Requirements for Data Quality

| Parameter | Method | Minimum Requirements |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Precision | Accuracy | Flow Accuracy | Completeness |
| Carbon Monoxide (CO) | Automated | 15% | 15% |  | 75% Quarterly |
| Ozone (O3) | Automated | 15% | 15% |  | 75% Seasonally |
| Nitrogen Oxides (NOX) | Automated | 15% | 15% |  | 75% Quarterly |
| Sulfur Dioxide (SO2) | Automated | 15% | 15% |  | 75% Quarterly |
| Total Suspended Particulate (TSP) | Manual | 20-80 ug/m3 5 ug/m3 >80 ug/m3 7% |  | 10% | 75% Quarterly |
| Inhalable Particulate (PM10) | Manual | 20-80 ug/m3 5 ug/m3 >80 ug/m3 7% |  | 7% | 75% Quarterly |
| PM10 | Automated | N/A |  | 7% | 75% Quarterly |
| Respirable Particulate (PM2.5) | ## Manual |  10% |  |  4% | 75% Quarterly |
| PM2.5 | Automated | ## N/A |  | ## N/A | 75% Quarterly |
| Lead | Manual |  10% | 10% | 10% | 75% Quarterly |

Table 2 - APCD Data Quality Objectives

| Parameter | Method | Objectives |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | Precision | Accuracy | Flow Accuracy | Completeness |
| Carbon Monoxide (CO) | Automated | ±10% | ±10% |  | 90% / annually |
| Ozone (03) | Automated | ±10% | ±10% |  | 90% / seasonally |
| Nitrogen Oxides (NOx) | Automated | ±10% | ±10% |  | 90% / annually |
| Sulfur Dioxide (SO2) | Automated | ±10% | ±10% |  | 90% / annually |
| Total Suspended Particulate (TSP) | Manual | 20-80 ug/m3 5 80 ug/m3 7% |  | 10% | 75% / quarterly |
| Inhalable Particulate (PM10) | Manual | 20-80 ug/m3 5 80 ug/m3 7% |  | 7% | 75% / quarterly |
|  | Automated | N/A |  | 4% total 6% main | 90% / quarterly |
| Respirable Particulate (PM2.5) | Manual | 10% |  | 4% | 90% / quarterly |
|  | Automated | N/A |  | 4% total 6% main | 90% / quarterly |
| Lead (Pb) | Manual | 10% | 10% | ` | 90% / quarterly |
| Wind Speed | Automated | N/A | 0.2 m/s and 5% |  | 90% / annually |
| Wind Direction | Automated | N/A | 5o |  | 90% / annually |
| Temperature | Automated | N/A | 0.5o C |  | 90% / annually |
| Gravimetric | Manual | 0.2 mg | 0.2 mg |  | 90% / annually |

1.0 STATUS OF THE MONITORING NETWORK

Most of the tables in this document have abbreviated site information.  More detail about sites, such as site I.D., county, and address can be found in Appendix A.  If more specific information is desired, complete and detailed site information can be found in the APCD Annual Network Review.

An important quality assurance concern is the degree of completeness of the data collected by the monitoring network.  In order for ambient monitoring data to be considered representative, EPA requires that valid data must be available for at least 75% of the averaging period.  The APCD has an internal goal of 90% data capture rate for each automated analyzer and 75% for each manual sampler.  For automated sampling systems, gaseous pollutant and meteorological monitors, these data capture values are calculated by dividing the number of valid hourly samples by the total number of hours in the sampling period (generally a day, a quarter or a year).  This is a conservative estimate of data completeness, since it does not take into consideration times when the sampler is disabled for routine maintenance, calibration, and audits.   For manual methods the data completeness is calculated by dividing the number of valid daily samples at a site by the number of days that site is scheduled to sample during a period of time (generally a quarter or a year).

Table 3 presents the data recovery for all gaseous pollutant monitors in the APCD network.  The overall average data recovery for gaseous pollutants during 2001 was 92%.  This data recovery rate is excellent and is similar to the recovery rate calculated for the past few years.  26 of 32 gaseous sites met the APCD 90% data capture goal.

The data capture rates for the particulate monitoring networks and the lead network are presented in Tables 4a and 4b.  The TSP network had an overall data recovery of 93% and the lead network had an overall recovery of 85%.  The PM10 network had an average annual data recovery of 88%, and the PM2.5 network was at 93% data completeness.  The sites with these manual samplers operate every day, once every other day, once ever third day, or once every six days depending on the site.  For more specific information on sampling frequency of the particulate networks, please refer to the APCD Annual Network Review.

All but one TSP site achieved the minimum 75% data completeness in 2001.

More than half of the PM10 sites obtained data recovery above 90% for 2001.  Most of the PM10 sites complied with the 75% EPA data completeness requirement in 2001.  The following sites did not meet the minimum criteria: Castlerock, Vail, Colorado Springs-Meadowland, Rifle, Ft. Collins-courthous, Montrose, Aspen, and Breckenridge.

In 2001 all but six of the 23 monitors achieved the desired 90% data recovery rate in the PM2.5 network and all achieved the minimum EPA acceptance criteria of 75%.

Data capture rates for all parameters of the 19 APCD meteorological monitoring sites were above 90% in 2001 with the following exceptions:

Lamar					due to power failure and communications problems.

Telluride				due to power failure and communications problems.

Table 3 - 2001 Gaseous Data Recovery

| SITE / STATION NAME | CO | O3 | SO2 | NOX |
| --- | --- | --- | --- | --- |
| Arvada | 95% | 99% |  |  |
| Boulder - S. Boulder Creek |  | 97% |  |  |
| Boulder - YMCA | 98% |  |  |  |
| Chatfield Reservoir * |  | 98% |  |  |
| Colorado Springs - GLEN | 88% |  |  |  |
| Colorado Springs - Hwy 24 | 99% |  |  |  |
| C.S. - US Air Force Academy |  | 99% |  |  |
| Denver - CAMP | 98% |  | 92% | 87% |
| Denver - Carriage | 99% | 96% |  |  |
| Denver - NJH | 95% |  |  |  |
| Denver Firehouse #6 | 99% |  |  |  |
| Ft. Collins | 98% | 88% |  |  |
| Grand Junction - Stocker | 88% |  |  |  |
| Greeley | 97% | 99% |  |  |
| Highlands |  | 92% |  |  |
| Longmont | 91% |  |  |  |
| NREL |  | 98% |  |  |
| Rocky Flats X-1 (N) |  | 98% |  |  |
| Rocky Flats X-3 (SE) |  |  |  | 93% |
| Rocky Flats X-5 (W) |  |  |  | 70% |
| South Adams Pump Station |  |  | 92% |  |
| Welby | 98% | 95% | 94% | 86% |
| Welch |  | 97% |  |  |
| Parameter Average Data Recovery | 96% | 96% | 93% | 84% |
| NOTES: All monitors run continuosly for 2001, with following exceptions: |  |  |  |  |
| (*1) = Seasonal only |  |  |  |  |
| (*2) = New Site |  |  |  |  |
| (*3) = Site Discontinued |  |  |  |  |

Table 4A – 2001 Particulate and Lead Data Recovery

| SITE / STATION NAME | PM2.5 | PM10 | TSP | Pb |
| --- | --- | --- | --- | --- |
| Adams City (primary) (*1) | 94% | 95% | 83% | 80% |
| Adams City (collocated) (*1) | 100% |  |  |  |
| Globeville Clinicare |  |  | 94% | 78% |
| Commerce City (primary) (*3) | 93% | 97% | 72% | 72% |
| Commerce City (collocated) (*3) | 93% |  |  |  |
| Brighton (*2) |  | 83% |  |  |
| Welby |  | 90% |  |  |
| Alamosa |  | 86% |  |  |
| Arapahoe Community College | 96% |  |  |  |
| Pagosa Springs - Middle School (*3) | 100% | 96% |  |  |
| Pagosa Springs (*1) | 96% | 96% |  |  |
| Longmont (particulate) | 95% | 96% |  |  |
| Boulder Chamber of Commerce | 94% | 98% |  |  |
| Hygiene |  | 96% |  |  |
| Delta | 84% | 87% |  |  |
| Paonia (*1) |  |  |  |  |
| Hotchkiss (*1) |  |  |  |  |
| Denver - CAMP (primary) | 96% | 98% | 97% | 97% |
| Denver - CAMP (collocated) | 94% | 92% |  |  |
| Denver - NJH | 89% | 96% |  |  |
| Denver - Gates (primary) |  | 84% | 98% | 85% |
| Denver - Gates (collocated) |  | 97% | 97% | 80% |
| Denver Visitor Center (*2) |  | 95% |  |  |
| Denver - LARS |  | 93% |  |  |
| Castle Rock (*1) |  | 73% |  |  |
| Parker (*1) | 97% |  |  |  |
| Vail - WSD (*1) |  | 73% |  |  |
| Elbert | 84% |  |  |  |
| Colorado Springs - Meadowlands | 98% | 70% |  |  |
| Colorado Springs - RBD (primary) | 93% | 89% | 98% | 95% |
| Colorado Springs - RBD (collocated) | 97% | 97% |  |  |
| Canon City |  | 97% |  |  |
| Parachute |  | 83% |  |  |
| Rifle (*1) |  | 70% |  |  |
| Glenwood Springs (*1) |  | 100% |  |  |
| Crested Butte |  | 93% |  |  |
| Mt. Crested Butte | 95% | 95% |  |  |
| Gunnison |  | 81% |  |  |
| Rocky Flats X-1 (N) (primary) (*1) |  | 97% | 97% |  |
| Rocky Flats X-1 (N) (collocated) (*1) |  |  | 97% |  |
| Rocky Flats X-2 (NE) (primary) (*1) |  | 97% | 97% |  |
| Rocky Flats X-2 (NE) (collocated) (*1) |  | 93% |  |  |
| Rocky Flats X-3 (SE) (*1) |  | 100% | 97% |  |
| Rocky Flats X-4 (S) (*1) |  | 97% | 90% |  |
| Rocky Flats X-5 (W) (*1) |  | 83% | 83% |  |
| Leadville |  |  | 90% | 90% |
| Durango - Platform |  | 88% |  |  |
| Durango - Park School | 77% | 80% |  |  |
| Durango - Courthouse |  | 87% |  |  |
| Ft Collins - Courthouse (*1) |  | 58% |  |  |
| Ft. Collins - CSU | 97% | 88% |  |  |
| Grand Junction - Health Dept. (primary) | 86% | 97% | 100% |  |
| Grand Junction - Health Dept. (collocated) | 97% |  |  |  |
| Grand Junction - Stocker (*1) |  | 87% |  |  |
| Montrose (*1) |  | 64% |  |  |
| Olathe (*1) |  | 79% |  |  |
| Aspen |  | 73% |  |  |
| Lamar Power Plant |  | 89% |  |  |
| Lamar Municipal Bldg |  | 85% |  |  |
| Pueblo Public Works | 98% | 82% |  |  |
| Steamboat Springs | 79% | 95% |  |  |
| Telluride | 89% | 93% |  |  |
| Breckenridge |  | 64% |  |  |
| Silverthorne Rec Center |  | 76% |  |  |
| Cripple Creek |  | 79% |  |  |
| Greeley | 91% | 94% |  |  |
| Platteville | 95% |  |  |  |
| Parameter Average Data Recovery | 93% | 88% | 93% | 85% |
| NOTES: Specific site sampling schedules can be found in Appendix A |  |  |  |  |
| (*1) = Site removed during 2001 |  |  |  |  |
| (*2) = Shut down part of year for site maintenance |  |  |  |  |
| (*3) = New Site |  |  |  |  |

Table 4B – Data Recovery for Continuous Patuculate Samplers

| Site/Station Name | Responsible Organization | Instrument | Analysis | % Data Completeness | NOTES |
| --- | --- | --- | --- | --- | --- |
| Welby | CDPHE | TEOM | PM10 | 96% | New Sampler |
| Denver - CAMP | CDPHE | TEOM | PM10 | 98% |  |
| Denver - CAMP | CDPHE | TEOM | PM2.5 | 92% | New Sampler |
| Grand Junction | Mesa County Health Dept. | Beta Gauge | PM10 | 91% |  |
| Aspen | Pitkin County Health Dept. | TEOM | PM10 (daily) | 97% |  |
| Aspen | Pitkin County Health Dept. | same as above | PM10 (hourly)) | 96% |  |
| Telluride | San Miguel Cnty. Hlth. Dept. | Beta Guage | PM10 | 53% | APCD not responsible for operation of this sampler - only for data reporting |

2.0 GASEOUS MONITORING NETWORK QUALITY ASSURANCE

The Air Pollution Control Division routinely monitors four gaseous pollutants; carbon monoxide (CO), ozone (O3), sulfur dioxide (SO2), and nitrogen dioxide (NO2).  Quality assurance of this gaseous monitoring network involves a series of daily tests, maintenance activities, and record keeping, as well as several external assessment and control functions not performed on a daily basis.

In 2001, the gaseous monitoring network was involved in three non-daily programs; (1) regularly scheduled assessments and calibrations, (2) the National Air Monitoring Stations (NAMS) and State and Local Air Monitoring Stations (SLAMS) precision and accuracy program, and (3) EPA NPAP interlaboratory comparison studies.  The Quality Assurance Unit is responsible for recommending corrective actions when monitoring problems are discovered through these programs.  It is the responsibility of the Continuous Monitoring Data Support Services (CMDSS) Unit to repair the monitors and perform the necessary corrective actions.

2.1 Assessment and Accuracy Audits

APCD staff perform two types of gaseous analyzer performance audits; assessment audits and accuracy audits.  These audits challenge the analyzer with pollutant gases of known concentration within the range of the analyzer.

The APCD CMDSS Unit conducts frequent, regularly scheduled assessment audits.  These assessment audits provide a means of gauging the response of an analyzer before calibration, adjustments, or repairs are made.  Immediately before analyzer calibration, a CMDSS Unit staff person introduces a known gas concentration near the full-scale response of the instrument.  If the instrument response shows an error greater than  10%, a full multi-point assessment audit is conducted prior to any analyzer adjustment.  Assessment audits are also conducted when problems with the monitors are known or suspected.  Personnel and equipment used in assessment audits are the same as those used for calibrations.

Accuracy audits are conducted with equipment independent of that used for calibration by the APCD Quality Assurance (QA) Unit.  These randomly scheduled audits, using equipment and personnel different than those used in instrument calibration, meet EPA's definition of an accuracy audit.  The results of accuracy audits are submitted to the EPA's national air quality database, the Aerometric Information Retrieval System - Air Quality Subsystem (AIRS-AQS).

Because both types of audits use gas standards of very high quality, APCD management considers these assessment and accuracy audits to be equivalent.  In the event of an audit failure, data validation procedures are identical, regardless of whether the CMDSS Unit or QA Unit staff discovered the problem.

Table 5 show the number of Quality Assurance Unit accuracy audits and CMDSS Unit assessment audits or calibrations conducted in 2001 at each of the APCD monitoring sites.  Of the 194 accuracy and assessment audits conducted by APCD staff during 2001, 22 audits showed unacceptable errors at full scale.  The APCD acceptance criterion for both accuracy and assessment audits is an instrument response less than 10% error at analyzer full scale.  In those instances where a relative error greater than 10% at full scale was discovered, a brief description of corrective actions taken to remedy the situation can be found in the right hand column.  A more detailed overview of all audit failures and the corrective actions taken can be found is Section 6 of this QA Report.

Table 5 - 2001 Gaseous Internal Audits Summary

| SITE / STATION NAME | Parameter | Assessment | Accuracy | Failed | Corrective Actions |
| --- | --- | --- | --- | --- | --- |
| Adams City - Nigra Station (*2) | SO2 | 1 | 0 |  |  |
| Arvada | O3 | 5 | 2 |  |  |
|  | CO | 5 | 2 |  |  |
| Boulder - S. Boulder Creek | O3 | 5 | 1 |  |  |
| Boulder - YMCA | CO | 4 | 2 |  |  |
| Chatfield Reservoir (*1) | O3 | 2 | 2 |  |  |
| Colorado Springs - GLEN | CO | 3 | 2 |  | Maintenance and Calibration |
| Colorado Springs - Highway 24 | CO | 3 | 2 |  |  |
| C.S. - US Air Force Academy | O3 | 4 | 2 |  |  |
| Denver - CAMP | CO | 5 | 2 |  |  |
|  | SO2 | 5 | 2 |  |  |
|  | NOx | 7 | 2 | 4 | Maintenance and Calibration |
| Denver - Carriage | O3 | 5 | 2 |  |  |
|  | CO | 5 | 2 |  |  |
| Denver - NJH | CO | 3 | 1 |  |  |
| Denver Firehouse #6 | CO | 2 | 2 |  |  |
| Ft. Collins | O3 | 5 | 2 |  |  |
|  | CO | 4 | 2 |  |  |
| Grand Junction - Stocker | CO | 3 | 2 |  |  |
| Greely | O3 | 4 | 2 |  |  |
|  | CO | 3 | 2 |  |  |
| Highlands | O3 | 6 | 2 |  |  |
| Longmont | CO | 3 | 2 |  |  |
| NREL | O3 | 5 | 2 | 1 | Maintenance and Calibration |
| Rocky Flats X-1 (N) | O3 | 5 | 2 |  |  |
| Rocky Flats X-3 (SE) (*3) | NOx | 4 | 1 | 3 | Maintenance and Calibration |
| Rocky Flats X-5 (W) (*3) | NOx | 6 | 1 | 5 | Maintenance and Calibration |
| Welby | O3 | 8 | 2 | 2 | Maintenance and Calibration |
|  | CO | 3 | 2 |  |  |
|  | SO2 | 4 | 2 | 1 | Maintenance and Calibration |
|  | NOx | 5 | 2 | 4 | Maintenance and Calibration |
| Welch | O3 | 5 | 2 | 2 |  |
| TOTAL GASEOUS AUDITS | 32 Analyzers | 136 | 58 | 22 | 4% Failure Rate |
| SUBTOTALS: |  |  |  |  |  |
| O3 Audits | 12 Analyzers | 50 | 24 | 5 | 7% Failure Rate |
| CO Audits | 13 Analyzers | 54 | 24 | 0 | 0% Failure Rate |
| SO2 Audits | 3 Analyzers | 10 | 4 | 1 | 7% Failure Rate |
| NOx Audits | 4 Analyzers | 22 | 6 | 16 | 57% Failure Rate |
| NOTES: |  |  |  |  |  |
| (*1) = Seasonal only |  |  |  |  |  |
| (*2) = New Site |  |  |  |  |  |
| (3) = Site closed in 2001 |  |  |  |  |  |

The Quality Assurance Unit completed 58 gaseous accuracy audits in 2001.  APCD QA Unit maintained compliance with the EPA requirement of auditing each SLAMS and NAMS analyzer a minimum of once per year, with a minimum of 25% of each type of analyzer to audited each quarter.

During 2001 there were 13 continuously operated carbon monoxide monitoring sites in the APCD network.  There were no audit failures of a CO analyzer.  There were 12 continuously operated ozone monitoring sites in the APCD network.   There were 5 assessment or accuracy audit failures of an O3 analyzer in 2001.  There were two continuously operated SO2 analyzers with one audit failure.

Nitrogen dioxide monitoring was conducted at four sites in 2001.  There were sixteen assessment or accuracy audit failures.  This high number of failures is consistent with previous years.  The Nitrogen dioxide measurement principle is highly sensitive to small flow and vacuum changes.  The failures are usually found, as expected, soon after routine maintenance has occurred.  The high number of these failures is a reflection of the CMDSS staff keeping tight controls on these instruments, and a reflection of their quick response time when a problem is suspected.  Additional tests and comparisons were performed on the NOx network and NOx equipment and the APCD TSP group concluded that all converters should be replaced in these analyzers.

Table 6 summarizes the total number of QA gaseous accuracy audits performed annually since 1990.

Figure 1 - Number of Gaseous APCD Accuracy Audits (1990-2001)

2.2 Data Validation

APCD staff perform a number of routine data validation procedures to ensure the quality of the ambient data collected by the APCD's monitoring network.  Site operators regularly perform diagnostic checks on monitors, and compare the monitor data logger and chart recorder response.  Ambient data are telemetered to APCD offices and are reviewed daily for high values or anomalous readings.  Daily span and zero tests are plotted in control chart format and reviewed.  The control chart limits for span readings are consistent with the error calculations used to assess analyzer performance.  Rather than setting individual analyzer action and warning limits, a warning limit of 7% error and an action limit of 10% error is used for all gaseous analyzers in the network.  The CMDSS Unit updates these control charts daily in order to identify problem situations and to initiate a prompt response when monitors begin to drift out of acceptable performance range.

The ambient data are also reviewed in response to failed accuracy or assessment audits.  In response to an audit error greater than 15%, a data quality assessment is conducted including review of operator logs, maintenance records, and analyzer control charts to identify the cause of the audit failure.  Current APCD policy is to invalidate all ambient data from the last analyzer calibration, valid QC check, or known failure point up to the date that corrective action (usually analyzer maintenance and recalibration) is completed.

2.3 EPA Interlaboratory Studies

In 2001 the APCD took part in national interlaboratory studies for carbon monoxide (CO), ozone (O3), sulfur dioxide (SO2), and nitrogen oxide (NO2).  The interlaboratory comparisons are part of the National Performance Audit Program (NPAP) operated by the EPA.  Unfortunately, the NPA program invalidated all ozone audits due to problems with their auditing equipment.  These interlaboratory studies are designed to assess the compliance of agencies operating monitors in the State and Local Air Monitoring System (SLAMS) network.  This program provides a means of comparing the APCD's monitoring performance with that of other participating agencies.  Although an organization can use the interlaboratory program to identify out-of-control situations, the interlaboratory results cannot be used to adjust or invalidate data.

The interlaboratory studies involved challenging analyzers with audit concentrations generated from EPA equipment.  These audit concentrations or "assigned values" are unknown to the APCD personnel conducting the assessment.  The EPA target criterion for acceptable audit performance is that the participant's value for each concentration be within 15% of the assigned value.  Participants reporting values outside these limits are urged to determine the cause for their poor performance and to take corrective action.  Participants reporting results within 10 -15% of the assigned value are encouraged to evaluate their measurement systems in order to reduce the size of the difference.

The interlaboratory kit used in 2001 consisted of a multi-blend gas cylinder (CO, SO2, and NO), clean air scrubbers, and a critical orifice dilution kit to regulate pollutant gas flow, and an ozone generating ultraviolet light source.  The NPAP interlaboratory audits were conducted during the first quarter of 2001.  Table 7 summarizes the 2001 NPAP audit results.  This table shows EPA values, each individual APCD analyzer’s response values, and the percent difference between the EPA “true” and the APCD “indicated” value for each of the three audit levels.  The chart also has regression values generated from these data, the EPA’s calculated “mean absolute % error”, and APCD’s calculated % analyzer full-scale error.  All the ozone and SO2 results were within the APCD's  10% full scale error and the EPA 15% mean error criteria.  It was later determined that a problem existed with our NOx analyzers.  The problem was identified and corrected.

Table 6 - NPAP Interlaboratory Gaseous Audit Results

|  | Parameter | Date | NPAP | APCD | NPAP | APCD | % Error | NPAP | APCD | % Error | NPAP | APCD | % Error | Slope | Intercept | Mean | % Full |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Site / Station Name | (reporting | of | Actual | Indicated | Actual | Indicated | on High | Actual | Indicated | Middle | Actual | Indicated | on Low |  |  | ABS % | Scale |
|  | units) | Audit | Value | Value | Value | Value | Level | Value | Value | Level | Value | Value | Level |  |  | Error | Error |
| Ft. Collins | CO (ppm) | 3/20/2001 | 0.00 | 0.10 | 43.05 | 42.9 | -0.3% | 16.91 | 16.3 | -3.6% | 6.95 | 6.5 | -6.5% | 0.999 | -0.259 | 3.5% | -0.6% |
| CO Springs - Hwy 24 | CO (ppm) | 3/27/2001 | 0.00 | 0.30 | 43.05 | 44.2 | 2.7% | 16.91 | 17.1 | 1.1% | 6.95 | 6.9 | -0.7% | 1.001 | -0.344 | 1.5% | -0.6% |
| CAMP | CO (ppm) | 3/22/2001 | 0.00 | 0.00 | 43.05 | 42.5 | -1.3% | 16.91 | 16.3 | -3.6% | 6.95 | 6.6 | -5.0% | 0.990 | -0.205 | 3.3% | -1.4% |
| Denver Firehouse #6 | CO (ppm) | 3/27/2001 | 0.00 | 0.40 | 43.05 | 43.9 | 2.0% | 16.91 | 16.9 | -0.1% | 6.95 | 6.9 | -0.7% | 1.015 | 0.042 | 0.9% | 1.6% |
| CAMP | SO2 (ppb) | 3/22/2001 | 0.00 | 1.00 | 412.07 | 429 | 4.1% | 161.86 | 162 | 0.1% | 66.57 | 65 | -2.4% | 1.044 | -2.834 | 2.2% | 3.8% |
| CAMP | NO2 (ppb) | 3/22/2001 | 1.00 | -2.00 | 377.70 | 395.0 | 4.6% | 177.70 | 185.0 | 4.1% | 73.00 | 73.0 | 0.0% | 1.055 | -3.255 | 2.9% | 4.8% |
| Rocky Flats W (X-5) | NO2 (ppb) | 3/15/2001 | 1.00 | -0.50 | 377.70 | 293.7 | -22.2% | 177.70 | 140.3 | -21.0% | 73.00 | 58.4 | -20.0% | 0.779 | 0.414 | 21.1% | -22.0% |
| Rocky Flats SE (X-3) | NO2 (ppb) | 3/16/2001 | 1.00 | 5.50 | 377.70 | 308.9 | -18.2% | 177.70 | 153.5 | -13.6% | 73.00 | 64.4 | -11.7% | 0.806 | 6.279 | 14.5% | -18.1% |
| NOTES: All Ozone audits were invalidated by ManTech due to a faulty ozone generator in the NPAP auditing device. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| It should also be noted that all NO2 audits were conducted with the same fauly auditing device, but ManTech didn't feel it was necessary to invalidate these audits. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

2.4 Precision and Accuracy Program

The APCD is required to conduct a one-point precision test every two weeks for each NAMS and SLAMS continuous monitor.  This precision test involves introducing an analyte gas of known concentration, the "actual value," and determining an analyzer response or "indicated value."  During this precision test, the analyte gas must pass through all of the filters and components of the ambient sampling system.  The actual value of the analyte gas is determined shortly after the analyzer calibration.  This value must fall within the range of 8-10 ppm for CO, and 0.080-0.100 ppm for O3, SO2, and NO2.

Current APCD practice is to perform these precision tests weekly, using the automated system developed for span gas delivery and diluting the span gas to the appropriate precision test value.  In addition to these automated precision tests, station operators, calibration personnel and audit personnel perform manually initiated tests to assess the ability of an analyzer to repeatedly measure a known analyte gas at relatively low concentration, typical of what may be found in ambient air.  This testing provides an assessment of the repeatability of a measurement system.  The values determined for this precision test system are sent to the EPA Aerometric Information Retrieval System (AIRS) database within 90 days of the end of each calendar quarter.

The EPA goal for single analyzer precision is 15%.  The APCD goal is to have all precision within 10%.  The weekly-automated precision checks are reviewed frequently, and a CMDSS staff member is sent out to any site where the precision values exceed the 7% warning limits to investigate the problem.

The results of the accuracy audits performed by the APCD's Quality Assurance group are submitted to AIRS at the end of each calendar quarter.  As in the precision test program, "actual" and "indicated" values are reported.  The accuracy results are reported at up to four concentration levels.  Level 1 "actual values" must fall within the range of 3-8 ppm for CO, and 0.030-0.080 ppm for O3, SO2, and NO2.  Level 2 values are 15-20 ppm and 0.150-0.200 ppm, and Level 3 values are 35-45 ppm and 0.350-0.450 ppm.

The results of this precision and accuracy program provide information about the ability of the APCD's network to accurately measure ambient pollutant concentrations at various levels.  The precision and accuracy data submitted to AIRS are used to calculate precision probability limits, and accuracy probability limits at three audit levels.  These calculated probability values are presented in Tables 7 - 10.  The upper and lower probability limits indicate the range of percent difference from the "actual value" that would include 95% of the "indicated values."  About 5% of the results of an individual precision or accuracy test would exceed these limits.  Ideally, the probability ranges would be very small, and would be centered at zero.  This indicates that the ambient data collected by the analyzer are both precise and accurate.

Shown in Tables 7 – 10 are quarterly precision probability limits by site and by reporting organization (CDPHE), and annual precision summaries by site and reporting organization.  They also show quarterly and annual accuracy probability limits by reporting organization, but only for quarters where more than one audit was performed of that gaseous type.

Table 7 - Carbon Monoxide Precision and Accuracy Probability Limits

(95% Confidence Interval)

| Site/Station Name | Quarter | Precision Probability |  | Audit Probability Limits |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lower Limit | Upper Limit | Level 1 |  | Level 2 |  | Level 3 |  |
|  |  |  |  | Lower | Upper | Lower | Upper | Lower | Upper |
| 2001 |  |  |  |  |  |  |  |  |  |
| Arvada | 1 | -1% | 2% |  |  |  |  |  |  |
|  | 2 | -4% | 3% |  |  |  |  |  |  |
|  | 3 | -2% | 2% |  |  |  |  |  |  |
|  | 4 | -5% | -2% |  |  |  |  |  |  |
|  | Annual | -4% | 3% |  |  |  |  |  |  |
| Boulder - YMCA | 1 | -1% | 2% |  |  |  |  |  |  |
|  | 2 | -6% | 8% |  |  |  |  |  |  |
|  | 3 | -1% | 2% |  |  |  |  |  |  |
|  | 4 | 0% | 0% |  |  |  |  |  |  |
|  | Annual | -3% | 4% |  |  |  |  |  |  |
| CAMP | 1 | -1% | 3% |  |  |  |  |  |  |
|  | 2 | -4% | -1% |  |  |  |  |  |  |
|  | 3 | -3% | 7% |  |  |  |  |  |  |
|  | 4 | -2% | 3% |  |  |  |  |  |  |
|  | Annual | -4% | 5% |  |  |  |  |  |  |
| Colo. Spr. - GLEN | 1 | -3% | 2% |  |  |  |  |  |  |
|  | 2 | -2% | 1% |  |  |  |  |  |  |
|  | 3 | -6% | 6% |  |  |  |  |  |  |
|  | 4 | -1% | 11% |  |  |  |  |  |  |
|  | Annual | -5% | 7% |  |  |  |  |  |  |
| Colo. Spr. - Hwy 24 | 1 | -1% | 1% |  |  |  |  |  |  |
|  | 2 | -6% | 9% |  |  |  |  |  |  |
|  | 3 | -3% | 2% |  |  |  |  |  |  |
|  | 4 | -2% | 1% |  |  |  |  |  |  |
|  | Annual | -4% | 5% |  |  |  |  |  |  |
| Denver - Carriage | 1 | -8% | 9% |  |  |  |  |  |  |
|  | 2 | -5% | 2% |  |  |  |  |  |  |
|  | 3 | -2% | 2% |  |  |  |  |  |  |
|  | 4 | 0% | 3% |  |  |  |  |  |  |
|  | Annual | -5% | 5% |  |  |  |  |  |  |
| Denver - Firehouse #6 | 1 | -2% | 3% |  |  |  |  |  |  |
|  | 2 | 1% | 4% |  |  |  |  |  |  |
|  | 3 | -1% | 6% |  |  |  |  |  |  |
|  | 4 | -4% | 5% |  |  |  |  |  |  |
|  | Annual | -2% | 5% |  |  |  |  |  |  |
| Denver - NJH | 1 | -3% | 9% |  |  |  |  |  |  |
|  | 2 | -6% | 4% |  |  |  |  |  |  |
|  | 3 | -5% | 5% |  |  |  |  |  |  |
|  | 4 | -2% | 0% |  |  |  |  |  |  |
|  | Annual | -5% | 6% |  |  |  |  |  |  |
| Ft. Collins | 1 | -5% | 3% |  |  |  |  |  |  |
|  | 2 | -6% | 2% |  |  |  |  |  |  |
|  | 3 | -4% | 0% |  |  |  |  |  |  |
|  | 4 | -7% | 5% |  |  |  |  |  |  |
|  | Annual | -6% | 3% |  |  |  |  |  |  |
| Grand Junction | 1 | -4% | 2% |  |  |  |  |  |  |
|  | 2 | -9% | 4% |  |  |  |  |  |  |
|  | 3 | -7% | 2% |  |  |  |  |  |  |
|  | 4 | -3% | 0% |  |  |  |  |  |  |
|  | Annual | -6% | 3% |  |  |  |  |  |  |
| Greeley | 1 | -2% | 2% |  |  |  |  |  |  |
|  | 2 | -7% | 5% |  |  |  |  |  |  |
|  | 3 | -9% | 4% |  |  |  |  |  |  |
|  | 4 | -6% | 4% |  |  |  |  |  |  |
|  | Annual | -6% | 4% |  |  |  |  |  |  |
| Longmont | 1 | -7% | 4% |  |  |  |  |  |  |
|  | 2 | -2% | 2% |  |  |  |  |  |  |
|  | 3 | -5% | 4% |  |  |  |  |  |  |
|  | 4 | -1% | 2% |  |  |  |  |  |  |
|  | Annual | -4% | 3% |  |  |  |  |  |  |
| Welby | 1 | -3% | 1% |  |  |  |  |  |  |
|  | 2 | -3% | 0% |  |  |  |  |  |  |
|  | 3 | -2% | 1% |  |  |  |  |  |  |
|  | 4 | -1% | 2% |  |  |  |  |  |  |
|  | Annual | -3% | 1% |  |  |  |  |  |  |
| CDPHE | 1 | -4% | 4% | -4% | 2% | 1% | 4% | 1% | 7% |
| 2001 Summary | 2 | -6% | 5% | -6% | 1% | -3% | 1% | -2% | 2% |
|  | 3 | -5% | 5% | -15% | 8% | -7% | 0% | -6% | 4% |
|  | 4 | -5% | 5% | -10% | 5% | -5% | 6% | -3% | 9% |
|  | Annual | -5% | 5% | -9% | 5% | -5% | 5% | -4% | 8% |

Table 8 - Ozone Precision and Accuracy Probability Limits

(95% Confidence Interval)

| Site/Station Name |  | Quarter |  | Precision Probability |  |  | Audit Probability Limits |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Lower Limit |  | Upper Limit | Level 1 |  |  |  | Level 2 |  |  |  | Level 3 |  |  |  |  |
|  |  |  |  |  |  |  | Lower |  | Upper |  | Lower |  |  | Upper | Lower |  | Upper |  |  |
| 2001 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Arvada |  | 1 |  | -5% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -2% |  | 3% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -5% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -10% |  | 10% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -6% |  | 6% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Chatfield Res. |  | 2 |  | -5% |  | 6% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -1% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -3% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Co. Spr. - USAFA |  | 1 |  | -5% |  | 6% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -4% |  | 2% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -8% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -4% |  | 1% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -5% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Denver - Carriage |  | 1 |  | -7% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -7% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -9% |  | 7% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -4% |  | 3% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -7% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Ft. Collins |  | 1 |  | -4% |  | 16% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -4% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -7% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -7% |  | 9% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -7% |  | 10% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Greeley |  | 1 |  | -8% |  | 11% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -5% |  | 9% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -1% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -22% |  | 26% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -11% |  | 15% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Highlands |  | 1 |  | -6% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | -4% |  | -1% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -10% |  | 14% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -8% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -8% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NREL |  | 1 |  | 1% |  | 6% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 2 |  | 2% |  | 9% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 3 |  | -6% |  | 5% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | 4 |  | -3% |  | 2% |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  | Annual |  | -4% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |
| R.F. North | 1 |  | -6% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | -3% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | 1% |  | 10% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  | 1% |  | 7% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Annual |  | -4% |  | 9% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| S. Boulder Creek | 1 |  | -2% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | -1% |  | 1% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | -5% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  | -5% |  | 7% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Annual |  | -3% |  | 4% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Welby | 1 |  | -19% |  | 6% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | -2% |  | 8% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | -13% |  | 9% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  | -21% |  | 11% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Annual |  | -16% |  | 10% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Welch | 1 |  | 1% |  | 14% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 2 |  | -2% |  | 15% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 3 |  | -3% |  | 10% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | 4 |  | -5% |  | 18% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Annual |  | -2% |  | 14% |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| CDPHE | 1 |  | -9% |  | 11% |  |  | -9% |  | 2% |  | -5% | 4% |  |  | -7% |  | 6% |  |
| 2001 Summary | 2 |  | -5% |  | 8% |  |  | -5% |  | 3% |  | -1% | 7% |  |  | 0% |  | 8% |  |
|  | 3 |  | -7% |  | 8% |  |  | -9% |  | 5% |  | -6% | 7% |  |  | -4% |  | 7% |  |
|  | 4 |  | -11% |  | 12% |  |  | -17% |  | 11% |  | -2% | 10% |  |  | -5% |  | 6% |  |
|  | Annual |  | -8% |  | 10% |  |  | -11% |  | 6% |  | -4% | 8% |  |  | -4% |  | 7% |  |

Table 9 - Sulfur Dioxide Precision and Accuracy Probability Limits

(95% Confidence Interval)

| Site/Station Name | Quarter | # Precision Probability |  | # Audit Probability Limits |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lower Limit | Upper Limit | Level 1 |  | Level 2 |  | Level 3 |  |
|  |  |  |  | Lower | Upper | Lower | Upper | Lower | Upper |
| 2001 |  |  |  |  |  |  |  |  |  |
| CAMP | 1 | -5% | 2% |  |  |  |  |  |  |
|  | 2 | -6% | 1% |  |  |  |  |  |  |
|  | 3 | -12% | 6% |  |  |  |  |  |  |
|  | 4 | -6% | -2% |  |  |  |  |  |  |
|  | Annual | -8% | 3% |  |  |  |  |  |  |
| S. Adams - Niagra | 2 | -2% | 9% |  |  |  |  |  |  |
|  | 3 | -9% | 10% |  |  |  |  |  |  |
|  | 4 | -6% | 16% |  |  |  |  |  |  |
|  | Annual | -6% | 12% |  |  |  |  |  |  |
| Welby | 1 | -9% | 5% |  |  |  |  |  |  |
|  | 2 | -4% | 2% |  |  |  |  |  |  |
|  | 3 | -4% | 2% |  |  |  |  |  |  |
|  | 4 | -13% | 4% |  |  |  |  |  |  |
|  | Annual | -8% | 4% |  |  |  |  |  |  |
| CDPHE | 1 | -7% | 4% |  |  |  |  |  |  |
| 2001 Summary | 2 | -7% | 7% |  |  |  |  |  |  |
|  | 3 | -9% | 6% |  |  |  |  |  |  |
|  | 4 | -13% | 11% |  |  |  |  |  |  |
|  | Annual | -9% | 7% | -6% | 4% | -5% | 6% | -2% | 8% |
| Note: Quarterly accuracy probability limits are not available since only one audit was performed each quarter. |  |  |  |  |  |  |  |  |  |

Table 10 - Nitrogen Dioxide Precision and Accuracy Probability Limits

(95% Confidence Interval)

| Site/Station Name | Quarter | Precision Probability |  | # Audit Probability Limits |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Lower Limit | Upper Limit | # Level 1 |  | # Level 2 |  | Level 3 |  |
|  |  |  |  | Lower | Upper | Lower | Upper | Lower | Upper |
| 2001 |  |  |  |  |  |  |  |  |  |
| CAMP | 1 | -7% | 14% |  |  |  |  |  |  |
|  | 2 | -29% | 17% |  |  |  |  |  |  |
|  | 3 | -8% | 6% |  |  |  |  |  |  |
|  | 4 | -5% | 6% |  |  |  |  |  |  |
|  | Annual | -15% | 13% |  |  |  |  |  |  |
| R. F. SouthEast | 1 | -8% | 4% |  |  |  |  |  |  |
|  | 2 | -13% | 11% |  |  |  |  |  |  |
|  | Annual | -10% | 7% |  |  |  |  |  |  |
| R. F. West | 1 | -23% | -29% |  |  |  |  |  |  |
|  | 2 | -12% | 0% |  |  |  |  |  |  |
|  | Annual | -22% | -21% |  |  |  |  |  |  |
| Welby | 1 | -14% | 15% |  |  |  |  |  |  |
|  | 2 | -14% | 3% |  |  |  |  |  |  |
|  | 3 | -7% | 5% |  |  |  |  |  |  |
|  | 4 | -8% | 7% |  |  |  |  |  |  |
|  | Annual | -12% | 9% |  |  |  |  |  |  |
| CDPHE | 1 | -15% | 17% | -17% | 27% | -20% | 29% | -18% | 28% |
| 2001 Summary | 2 | -18% | 9% |  |  |  |  |  |  |
|  | 3 | -8% | 5% |  |  |  |  |  |  |
|  | 4 | -6% | 6% |  |  |  |  |  |  |
|  | Annual | -15% | 12% | -33% | 22% | -24% | 19% | -25% | 17% |
| Note: Accuracy probability limits available only if more than one audit was performed during the quarter. |  |  |  |  |  |  |  |  |  |

The annual CDPHE precision results for CO, O3, SO2, and NOx are graphed in Figures 2 – 5 along with the last 10 years of precision results so that so that network trends can be observed.

Figure 2 - APCD Annual Comparison of Carbon Monoxide Precision Probability Limits (1990 - 2001)

(95 % Confidence Interval)

Figure 3 - Figure 3 - APCD Annual Comparison of Ozone Precision Probability Limits (1990 - 2001)

(95% Confidence Interval)

Figure 4 - Figure 4 - APCD Annual Comparison of Sulfur Dioxide Precision Probability Limits (1990-2001)

(95% Confidence Interval)

Figure 5 - Figure 5 - APCD Annual Comparison of Nitrogen Dioxide Precision Probability Limits (1991 - 2001)

(95% Confidence Interval)

2.5 System Audits

A system audit is a qualitative inspection and review of a measurement system.  For the APCD's gas monitoring network, these audits offer an assessment of field operations at a given site.  Emphasis in these system audits is placed on safety and security at the site, and the adequacy of field procedures and record keeping.  APCD's gaseous monitoring stations received system audits in 2001.  The system audits revealed thorough maintenance and record keeping at all of the stations.  Very few problems were noted.  The APCD goal of conducting system audits at each of the gaseous monitoring stations was not met in 2001.  Eleven gaseous sites were audited.

2.6 Additional Audits

The APCD occasionally performs audits of gas monitoring networks operated by other public and private entities in Colorado.  Although not part of the APCD's monitoring network, the APCD has an interest in the quality of these data since they may be used in the determination of ambient background concentrations in Colorado.  Ozone audits were conducted at Rocky Mountain National Park and Mesa Verde National Park.  An SO2 audit was conducted at Buffalo Pass.

2.7 Certification of Audit Standards

For CO, NO2, and SO2 audits and calibrations, the APCD uses compressed gas cylinders whose accuracy is certified by the manufacturer.  These certifications are updated approximately every 18 months, and are conducted by the vendor using the Revised EPA Protocol For Assay and Certification of Compressed Gas Calibration Standards.  For SO2 and NO2 audit flow rates, the state uses NIST-traceable electronic flowmeters to measure dilution air flow.

The traceability of APCD ozone calibration and audit standards is maintained through the use of certified transfer standards.  The ozone audit device is certified regularly against the state's laboratory standard.  This standard is also employed in certifying the state's gaseous ozone calibrators, so that a common traceability is maintained for the APCD ozone monitoring network.  The laboratory standard is compared annually to a Standard Reference Photometer (SRP) operated by the EPA.

EPA Region VIII closed their Air Quality Assurance Laboratory in 1992.  The APCD was able to make arrangements to compare this laboratory standard against standard reference photometer located at the EPA Region VII office in Kansas City, Missouri.

Results of SRP comparisons made since 1992 are shown below.

Acceptable comparison results are a regression slope of 1.00 0.03 and a regression intercept of less than 5 ppb.  The APCD's laboratory standard has consistently met these criteria.

Date           		Equation

February 18, 2000	State Standard = 0.995 * SRP + 1.0 ppb (EPA VII)

March 13, 1997   	State Standard = 1.022 * SRP + 1.4 ppb (EPA VII)

March 13, 1996   	State Standard = 1.008 * SRP + 0.7 ppb (EPA VII)

March 16, 1995   	State Standard = 0.999 * SRP + 3.1 ppb (EPA VII)

March 17, 1994   	State Standard = 1.011 * SRP + 3.5 ppb (EPA VII)

March 17, 1993   	State Standard = 0.990 * SRP + 1.5 ppb (EPA VII)

Sept. 29, 1992   	State Standard = 0.978 * SRP + 2.0 ppb (EPA VII)

3.0 PARTICULATE MONITORING NETWORK QUALITY ASSURANCE

The Air Pollution Control Division (APCD) routinely monitors three size classes of particulate matter.  Total suspended particulates (TSP) includes particulate matter with a diameter of about 30 microns or less.  Inhalable particulate matter (PM10) includes all particulate matter with a diameter of 10 microns or smaller.  Respirable particulate matter (PM2.5)includes all particulate matter with a diameter of 2.5 microns or less.  Filters collected at some sites in the TSP network are also chemically analyzed for lead (Pb), arsenic (As), cadmium (Cd), and sulfate (SO4=).

The particulate networks quality assurance functions are centered around three activities; (1) routine field audits, (2) the National Air Monitoring Stations (NAMS) and State and Local Air Monitoring Stations (SLAMS) precision and accuracy program, and (3) EPA national NPAP and PEP programs.

3.1 Accuracy Audits

The particulate sampler flow audits conducted by the APCD are evaluated using two criteria.  The first criteria is that sampler flow at the current operating set point, measured using an independent flow standard, must be within 7% of the flow determined at the time of calibration for PM10 and TSP samplers.  For a PM2.5 sampler,  the instrument flow rate at the time of audit, measured using an independent flow transfer standard, must be within 4%.  The second criteria has to do with instrument design flow rate.  For PM10 samplers, the sampler flow rate (corrected to conditions of standard temperature and pressure) must also be within 10% of the nominal sampler design flow air rate of 1.132 cubic meters per minute.  For PM2.5 instruments, the actual sampling flow rate must be within 5% of the optimal instrument design flow conditions of 16.67L/min calculated in using actual local conditions.

Field audits of sampler flow rates were performed at all operating TSP, PM10, and PM2.5 sites in Colorado during 2001.  PM2.5 sites were audited quarterly.  PM10 and TSP sites were audited at least semi-annually at all sites, and quarterly at most of the sites.  Audit results for all particulate sites are summarized in Table 11.  Corrective actions (which included sampler maintenance, recalibration, and data invalidation) were performed after analyzing the results of the out-of-specification audits.  Summaries of these corrective actions can be found in Section 6 of this 2001 QA Report.

According to the 1997 revised CFR Part 58 Appendix A additional calculations to determine bias of the PM2.5 independent accuracy audits are required.  Table 12 summarizes the bias on PM2.5 QA accuracy audits.

Table 11 - 2001 Particulate Internal Audits Summary

|  |  | Number of Audits |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Site / Station Name | PM2.5 | TSP | PM10 | Failures | Corrective Actions |
| Adams City (primary) (*1) | 2 | 2 | 8 |  |  |
| Adams City (collocated) (*1) | 2 |  |  |  |  |
| Commerce City (primary) (*4) | 4 | 4 | 16 | 3 | PM2.5, TSP, see CORR ACT chart |
| Commerce City (collocated) (*4) | 3 |  |  |  |  |
| Globeville Clinicare |  | 6 |  |  |  |
| Brighton (*2) |  |  | 6 |  |  |
| Welby |  |  | 4 |  |  |
| Alamosa |  |  | 24 | 2 | MFC adjustment |
| Arapahoe Community College | 5 |  |  |  |  |
| Pagosa Springs (*1) | 2 |  | 12 | 2 | MFC adjustment |
| Pagosa Spr. Middle School (*4) | 2 |  | 12 |  |  |
| Longmont | 4 |  | 8 |  |  |
| Boulder Chamber of Commerce | 4 |  | 4 |  |  |
| Hygiene |  |  | 16 | 2 | Warnings |
| Delta | 5 |  | 6 |  |  |
| Paonia (*1) |  |  | 1 |  |  |
| Hotchkiss (*1) |  |  | 1 |  |  |
| Denver - CAMP (primary) | 6 | 4 | 5 | 2 | Warnings |
| Denver - CAMP (collocated) (*3) | 2 |  | 4 |  |  |
| Denver - Gates (primary) |  | 4 | 4 |  |  |
| Denver - Gates (collocated) |  | 4 | 4 |  |  |
| Denver Visitor Center |  |  | 16 | 2 | Warnings |
| Denver - LARS |  |  | 6 |  |  |
| Castle Rock (*1) |  |  | 2 |  |  |
| Parker (*1) | 5 |  |  | 1 | Partial/seals replaced/press fail |
| Vail - WSD (*1) |  |  | 2 |  |  |
| Elbert | 4 |  |  |  |  |
| Colorado Springs - Meadowlands | 4 |  | 8 |  |  |
| Colorado Springs - RBD (primary) | 4 | 4 | 4 |  |  |
| Co. Springs - RBD (collocated) (*5) | 3 |  | 4 | 1 | Warning |
| Canon City |  |  | 4 |  |  |
| Parachute |  |  | 6 |  |  |
| Rifle (*1) |  |  | 2 |  |  |
| Glenwood Springs (*1) |  |  | 2 |  |  |
| Crested Butte |  |  | 8 |  |  |
| Mt. Crested Butte | 4 |  | 16 |  |  |
| Gunnison |  |  | 8 |  |  |
| Rocky Flats X-1 (N) (primary) (*1) |  | 2 | 2 |  |  |
| Rocky Flats X-1 (N) (collocated) (*1) |  | 2 |  |  |  |
| Rocky Flats X-2 (NE) (primary) (*1) |  | 2 | 2 |  |  |
| Rocky Flats X-2 (NE)(collocated) (*1) |  |  | 2 |  |  |
| Rocky Flats X-3 (SE) (*1) |  | 2 | 2 |  |  |
| Rocky Flats X-4 (S) (*1) |  | 2 | 2 |  |  |
| Rocky Flats X-5 (W) (*1) |  | 2 | 2 |  |  |
| Leadville |  | 4 |  |  |  |
| Durango - Platform |  |  | 24 | 3 | MFC adjustment |
| Durango - Park School | 4 |  | 12 | 4 | PM2.5, PM10, see CORR ACT chart |
| Durango - Courthouse |  |  | 12 | 3 | Warning/MFC adjustment |
| Ft Collins - Courthouse |  |  | 6 | 1 | Delete data/replace motor/calibrate |
| Ft. Collins - CSU | 4 |  | 8 | 1 | Warning |
| G.J. - Health Dept.(primary) (*6) | 4 | 4 | 4 |  |  |
| G.J. - Health Dept.(collocated) | 4 |  |  |  |  |
| Grand Junction - Stocker (*1) |  |  | 2 |  |  |
| Grand Junction - Traffic (*4) |  | 8 |  | 2 | Warnings |
| Montrose |  |  | 5 |  |  |
| Olathe |  |  | 8 |  |  |
| Aspen |  |  | 8 |  |  |
| Lamar Power Plant |  |  | 24 | 1 | MFC adjustment |
| Lamar Municipal Bldg |  |  | 24 | 2 | MFC adjustment |
| Pueblo Public Works | 4 |  | 8 |  |  |
| Steamboat Springs | 4 |  | 20 |  |  |
| Telluride | 4 |  | 12 |  |  |
| Breckenridge |  |  | 10 | 1 | MFC adjustment |
| Silverthorne |  |  | 3 |  |  |
| Cripple Creek |  |  | 16 | 3 | Warning |
| Greeley | 4 |  | 8 |  |  |
| Platteville | 4 |  |  | 1 | Partial /temperature probe calibrated |
| Total Particulate Audits | 101 | 56 | 449 | 37 | See Corrective Action (Section 6) |
| # of Failures | 4 partial | 5 | 28 |  |  |
| % Failure Rate | 4% | 9% | 6% |  |  |
| Number of Analyzers in Network | 26 | 17 | 115 |  |  |
| NOTES: Sampling schedule summaries can be found in Table 1 More detailed site, monitor, and sampling schedule information can be found in the Annual Network Review. Complete CORRECTIVE ACTION list can be found in Table 21 (*1) Discontinued operation during 2001 (*2) Site shut down part of year for site maintenance (*3) PM 2.5 monitor started operation during 2001 (*4) New Site (*5) Collocated PM2.5 instrument removed (*6) New TSP site |  |  |  |  |  |

| Site/ Station Name | Method Code | QUARTER | Accuracy Bias |
| --- | --- | --- | --- |
| Adams City Primary | 118 | Annual | -1.23% |
| Adams City Collocated | 118 | Annual | -0.65% |
| Commerce City Primary | 118 | Annual | 0.91% |
| Commerce City Collocated | 118 | Annual | 0.15% |
| Arapahoe Community College | 118 | Annual | -1.03% |
| Pagosa Springs | 117 | Annual | 0.26% |
| Pagosa Springs School | 117 | Annual | -0.51% |
| Longmont | 118 | Annual | 1.86% |
| Boulder | 118 | Annual | 0.10% |
| Delta | 117 | Annual | -1.00% |
| CAMP | 118 | Annual | -0.08% |
| Parker | 118 | Annual | 0.83% |
| Elbert County | 118 | Annual | 0.42% |
| Meadowlands | 118 | Annual | -0.05% |
| RBD Primary | 118 | Annual | 0.26% |
| RBD Collocated | 118 | Annual | -1.04% |
| Mt. Crested Butte | 117 | Annual | -1.65% |
| Durango | 117 | Annual | -0.38% |
| Ft. Collins | 118 | Annual | -0.68% |
| Grand Junction Primary | 118 | Annual | -0.93% |
| Grand Junction Collocated | 118 | Annual | 0.03% |
| Pueblo | 118 | Annual | 0.02% |
| Steamboat Springs | 117 | Annual | 0.26% |
| Telluride | 117 | Annual | -1.08% |
| Greeley | 118 | Annual | -0.57% |
| Platteville | 118 | Annual | 0.09% |
| CDPHE | 117 | 1 | -0.97% |
| 2001 Summary | 117 | 2 | -0.67% |
|  | 117 | 3 | -1.07% |
|  | 117 | 4 | -0.02% |
|  | 117 | Annual | -0.68% |
|  | 118 | 1 | 0.15% |
|  | 118 | 2 | -0.96% |
|  | 118 | 3 | -1.46% |
|  | 118 | 4 | 1.89% |
|  | 118 | Annual | -0.14% |
|  | Both | 1 | -0.14% |
|  | Both | 2 | -0.89% |
|  | Both | 3 | -1.37% |
|  | Both | 4 | 1.41% |
|  | Both | Annual | -0.27% |

Table 12 – Bias on Accuracy Audits for PM2.5

Figure 6 shows trends in the number of QA particulate accuracy audits performed by APCD QA staff over the past 12 years.  Typically, well over 400 particulate audits are performed each year.  In 2001, over 600 particulate audits were performed.

Figure 6 - Number of APCD Particulate Accuracy Audits Performed Annually (1990 - 2001)

Figure 6 – Particulate Accurac y Audits (1990 – 2000)

3.2 Data Validation

The APCD performs a number of routine data validation procedures to ensure the quality of the ambient data collected by the particulate monitoring network.  Site operators regularly perform flow checks on monitors before and after each sampling interval to ensure that they are operating within acceptable limits.  During data processing, the CDPHE Laboratory Division ensures that indicated sampler flow rates are within range, if the flow is out of range the sample is voided.  Ambient data are also reviewed in response to failed accuracy audits   10%.

3.3 EPA Interlaboratory Studies

In 2001, the APCD participated in the national interlaboratory comparison study for PM10  and TSP measurements.  These interlaboratory comparisons are part of the National Performance Audit Program (NPAP) operated by the EPA.  This program provides a means of comparing the APCD's monitoring performance with that of other agencies.  Although an organization can use the interlaboratory program to identify out-of-control situations, the interlaboratory results cannot be used to adjust or invalidate data.

The interlaboratory kit used in 2001consisted of a series of six orifice plates.  Five of the plates are used with TSP sampler audits, a single orifice plate is used for PM10 audits.  The current sampler flow rate, as determined from the most recent calibration, is compared with a flow rate determined using the EPA audit orifice.  The EPA target criterion for acceptable audit performance is that the flow difference be within 15% of the EPA value.  Participants reporting values outside these limits are urged to determine the cause for their poor performance and to take corrective action.  Participants reporting results within 10 -15% of the assigned value are encouraged to evaluate their measurement systems in order to reduce the size of the difference.

In 1993 the NPAP program requested that the APCD and other reporting organizations target specific sampling sites which had monitored PM10 concentrations approaching or exceeding that of the National Ambient Air Quality Standard.   In order to comply with this request, APCD performed 92 PM10 interlaboratory comparisons in 2001.   This large number of audits occurred because the agency responsible for sending out the auditing devices did not send out the 2000 kit until the first quarter of 2001.  Therefore, there were no audits in 2000.  One TSP interlaboratory comparison was completed in 2001.  The results of the 2001 comparisons are presented in Table 13.  This table shows that all but two of the APCD samplers met the EPA target for flow differences.  The average of the absolute flow differences of all the PM10 audits was within 3% of EPA’s “true” value.

Table 13 - NPAP Interlab TSP and PM10 Audit Results

| Audit | Audit |  | Site | Instrument | Plate | Reported | Actual | % |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Parameter | Date |  | Name | Identification | Number | Value | Value | Difference |
| PM10 | 11/20/2001 |  | Commerce City | 21972 | 88 | 1.103 | 1.091 | 1.1% |
| PM10 | 11/20/2001 |  | Commerce City | 31963 | 88 | 1.188 | 1.166 | 1.9% |
| PM10 | 11/20/2001 |  | Commerce City | 11975 | 88 | 1.158 | 1.141 | 1.5% |
| PM10 | 11/20/2001 |  | Commerce City | 17406 | 88 | 1.159 | 1.141 | 1.6% |
| PM10 | 3/15/2001 |  | Alamosa | 1 | 88 | 1.130 | 1.144 | -1.2% |
| PM10 | 3/15/2001 |  | Alamosa | 2 | 88 | 1.179 | 1.200 | -1.7% |
| PM10 | 3/15/2001 |  | Alamosa | 3 | 88 | 1.167 | 1.119 | 4.3% |
| PM10 | 3/15/2001 |  | Alamosa | 4 | 88 | 1.165 | 1.124 | 3.7% |
| PM10 | 11/27/2001 |  | Alamosa | 6000 | 88 | 1.139 | 1.160 | -1.8% |
| PM10 | 11/27/2001 |  | Alamosa | 4975 | 88 | 1.150 | 1.173 | -2.0% |
| PM10 | 11/27/2001 |  | Alamosa | 4070 | 88 | 0.888 | 0.914 | -2.8% |
| PM10 | 11/27/2001 |  | Alamosa | 12770 | 88 | 1.244 | 1.246 | -2.0% |
| PM10 | 3/15/2001 |  | Pagosa Springs | 1 | 88 | 1.201 | 1.056 | 13.7% |
| PM10 | 3/15/2001 |  | Pagosa Springs | 2 | 88 | 1.101 | 1.043 | 5.6% |
| PM10 | 3/15/2001 |  | Pagosa Springs | 3 | 88 | 1.127 | 1.005 | 12.1% |
| PM10 | 3/15/2001 |  | Pagosa Springs | 4 | 88 | 1.096 | 1.064 | 3.0% |
| PM10 | 11/28/2001 |  | Pagosa Springs | 5034 | 88 | 1.101 | 1.099 | 0.2% |
| PM10 | 11/28/2001 |  | Pagosa Springs | 6020 | 88 | 1.141 | 1.160 | -1.7% |
| PM10 | 11/28/2001 |  | Pagosa Springs | 17370 | 88 | 1.176 | 1.177 | -0.1% |
| PM10 | 11/28/2001 |  | Pagosa Springs | 5033 | 88 | 1.100 | 1.118 | -1.7% |
| PM10 | 3/8/2001 |  | Delta | 1 | 88 | 1.119 | 1.059 | 5.6% |
| PM10 | 11/8/2001 |  | Delta | 1 | 88 | 1.104 | 1.092 | 1.1% |
| PM10 | 3/12/2001 |  | Canon City | 1 | 88 | 1.184 | 1.144 | 3.5% |
| PM10 | 11/27/2001 |  | Canon City | 1 | 88 | 1.113 | 1.151 | -3.3% |
| PM10 | 3/13/2001 |  | Crested Butte | 1 | 88 | 1.110 | 1.141 | -2.8% |
| PM10 | 3/13/2001 |  | Crested Butte | 2 | 88 | 1.114 | 1.126 | -1.1% |
| PM10 | 11/5/2001 |  | Crested Butte | 5031 | 88 | 1.163 | 1.196 | -2.7% |
| PM10 | 11/5/2001 |  | Crested Butte | 15861 | 88 | 1.153 | 1.144 | 0.8% |
| PM10 | 3/14/2001 |  | Mt. Crested Butte | 1 | 88 | 1.122 | 1.132 | -0.9% |
| PM10 | 3/14/2001 |  | Mt. Crested Butte | 2 | 88 | 1.160 | 1.195 | -2.9% |
| PM10 | 3/14/2001 |  | Mt. Crested Butte | 3 | 88 | 1.159 | 1.159 | 0.0% |
| PM10 | 3/14/2001 |  | Mt. Crested Butte | 4 | 88 | 1.159 | 1.189 | -2.5% |
| PM10 | 11/5/2001 |  | Mt. Crested Butte | 2570 | 88 | 1.154 | 1.172 | -1.5% |
| PM10 | 11/5/2001 |  | Mt. Crested Butte | 5040 | 88 | 1.139 | 1.151 | -1.0% |
| PM10 | 11/5/2001 |  | Mt. Crested Butte | 6070 | 88 | 1.179 | 1.190 | -0.9% |
| PM10 | 11/5/2001 |  | Mt. Crested Butte | 17762 | 88 | 1.116 | 1.146 | -2.6% |
| PM10 | 3/14/2001 |  | Durango Platform | 4 | 88 | 1.133 | 1.162 | -2.5% |
| PM10 | 3/14/2001 |  | Durango Platform | 3 | 88 | 1.157 | 1.195 | -3.2% |
| PM10 | 3/14/2001 |  | Durango Platform | 2 | 88 | 1.104 | 1.178 | -6.3% |
| PM10 | 3/14/2001 |  | Durango Platform | 1 | 88 | 1.190 | 1.220 | -2.5% |
| PM10 | 11/29/2001 |  | Durango Platform | 17475 | 88 | 1.126 | 1.130 | -0.4% |
| PM10 | 11/29/2001 |  | Durango Platform | 16126 | 88 | 1.122 | 1.122 | 0.0% |
| PM10 | 11/29/2001 |  | Durango Platform | 17358 | 88 | 1.166 | 1.159 | 0.6% |
| PM10 | 11/29/2001 |  | Durango Platform | 5036 | 88 | 1.126 | 1.179 | -4.5% |
| PM10 | 3/14/2001 |  | Durango School | 2 | 88 | 1.244 | 1.274 | -2.4% |
| PM10 | 3/14/2001 |  | Durango School | 1 | 88 | 1.154 | 1.171 | -1.5% |
| PM10 | 11/28/2001 |  | Durango School | 31605 | 88 | 1.148 | 1.156 | -0.7% |
| PM10 | 11/28/2001 |  | Durango School | 8856 | 88 | 1.001 | 1.006 | -0.5% |
| PM10 | 3/14/2001 |  | Durango Courthouse | 2 | 88 | 1.094 | 1.009 | 8.4% |
| PM10 | 3/14/2001 |  | Durango Courthouse | 1 | 88 | 1.118 | 1.133 | -1.3% |
| PM10 | 11/28/2001 |  | Durango Courthouse | 10905 | 88 | 1.023 | 1.029 | -0.6% |
| PM10 | 11/28/2001 |  | Durango Courthouse | 17352 | 88 | 1.105 | 1.119 | -1.2% |
| PM10 | 3/8/2001 |  | Montrose | 1 | 88 | 1.011 | 0.861 | 17.4% |
| PM10 | 11/8/2001 |  | Montrose | 1 | 88 | 1.110 | 1.118 | -0.7% |
| PM10 | 3/5/2001 |  | Lamar Power Plant | 1 | 88 | 1.155 | 1.181 | -2.2% |
| PM10 | 3/5/2001 |  | Lamar Power Plant | 2 | 88 | 1.204 | 1.232 | -2.2% |
| PM10 | 3/5/2001 |  | Lamar Power Plant | 3 | 88 | 1.186 | 1.222 | -3.0% |
| PM10 | 3/5/2001 |  | Lamar Power Plant | 4 | 88 | 1.193 | 1.230 | -3.0% |
| PM10 | 10/30/2001 |  | Lamar Power Plant | 2069 | 88 | 1.157 | 1.162 | -0.4% |
| PM10 | 10/30/2001 |  | Lamar Power Plant | 11870 | 88 | 1.216 | 1.219 | -0.3% |
| PM10 | 10/30/2001 |  | Lamar Power Plant | 19142 | 88 | 1.238 | 1.244 | -0.5% |
| PM10 | 3/5/2001 |  | Lamar Muni Blgd | 4 | 88 | 1.173 | 1.204 | -2.6% |
| PM10 | 3/5/2001 |  | Lamar Muni Blgd | 3 | 88 | 1.144 | 1.177 | -2.8% |
| PM10 | 3/5/2001 |  | Lamar Muni Blgd | 2 | 88 | 1.136 | 1.182 | -3.9% |
| PM10 | 3/5/2001 |  | Lamar Muni Blgd | 1 | 88 | 1.204 | 1.242 | -3.0% |
| PM10 | 10/30/2001 |  | Lamar Muni Blgd | 21277 | 88 | 1.081 | 1.073 | 0.7% |
| PM10 | 10/30/2001 |  | Lamar Muni Blgd | 6030 | 88 | 1.141 | 1.136 | 0.5% |
| PM10 | 10/30/2001 |  | Lamar Muni Blgd | 6050 | 88 | 1.121 | 1.116 | 0.4% |
| PM10 | 10/30/2001 |  | Lamar Muni Blgd | 17386 | 88 | 1.095 | 1.087 | 0.7% |
| PM10 | 3/19/2001 |  | Steamboat Springs | 1 | 88 | 1.163 | 1.190 | -2.2% |
| PM10 | 3/19/2001 |  | Steamboat Springs | 2 | 88 | 1.189 | 1.220 | -2.5% |
| PM10 | 3/19/2001 |  | Steamboat Springs | 3 | 88 | 1.142 | 1.187 | -3.8% |
| PM10 | 3/19/2001 |  | Steamboat Springs | 4 | 88 | 1.183 | 1.218 | -2.9% |
| PM10 | 11/13/2001 |  | Steamboat Springs | 11222 | 88 | 1.127 | 1.154 | -2.4% |
| PM10 | 11/13/2001 |  | Steamboat Springs | 6080 | 88 | 1.165 | 1.186 | -1.8% |
| PM10 | 11/13/2001 |  | Steamboat Springs | 14384 | 88 | 1.137 | 1.161 | -2.0% |
| PM10 | 11/13/2001 |  | Steamboat Springs | 19035 | 88 | 1.097 | 1.107 | -0.9% |
| PM10 | 3/13/2001 |  | Telluride | 2 | 88 | 1.145 | 1.177 | -2.7% |
| PM10 | 3/13/2001 |  | Telluride | 1 | 88 | 1.092 | 1.129 | -3.3% |
| PM10 | 11/8/2001 |  | Telluride | 17320 | 88 | 1.125 | 1.119 | 0.6% |
| PM10 | 11/8/2001 |  | Telluride | 21953 | 88 | 1.167 | 1.147 | 1.7% |
| PM10 | 3/26/2001 |  | CAMP | primary | 88 | 1.129 | 1.091 | 3.5% |
| PM10 | 3/26/2001 |  | CAMP | collocated | 88 | 1.119 | 1.163 | -3.8% |
| PM10 | 3/9/2001 |  | Vail | 1 | 88 | 1.139 | 1.216 | -6.3% |
| PM10 | 3/8/2001 |  | Stocker - GJ | 1 | 88 | 1.009 | 1.037 | -2.7% |
| PM10 | 3/8/2001 |  | Grand Junction - Health | 1 | 88 | 1.150 | 1.162 | -1.0% |
| PM10 | 3/9/2001 |  | Glenwood | 1 | 88 | 1.131 | 1.168 | -3.2% |
| PM10 | 3/9/2001 |  | Aspen | 1 | 88 | 1.173 | 1.176 | -0.2% |
| PM10 | 3/9/2001 |  | Aspen | 2 | 88 | 0.971 | 0.614 | 58.1% |
| PM10 | 3/9/2001 |  | Rifle | 1 | 88 | 1.131 | 1.154 | -2.0% |
| PM10 | 3/8/2001 |  | Olathe | 1 | 88 | 1.131 | 1.117 | 1.3% |
| PM10 | 3/8/2001 |  | Olathe | 2 | 88 | 1.156 | 1.107 | 4.4% |
| Average % Difference |  |  |  |  |  |  |  | 0.3% |
| Average Absolute % Difference |  |  |  |  |  |  |  | 3.2% |
| Range |  |  |  |  |  |  | -6.3% to 58.1% |  |
| TSP | 11/20/2001 |  | Commerce City | primary | 5 | 0.524 | 0.579 | -12.3% |
| TSP | 11/20/2001 |  | Commerce City | primary | 7 | 0.641 | 0.716 | -10.5% |
| TSP | 11/20/2001 |  | Commerce City | primary | 10 | 0.778 | 0.822 | -5.4% |
| TSP | 11/20/2001 |  | Commerce City | primary | 13 | 0.838 | 0.880 | -4.8% |
| TSP | 11/20/2001 |  | Commerce City | primary | 18 | 0.910 | 0.946 | -3.8% |
| Average % Difference |  |  |  |  |  |  |  | -7.4% |
| Average Absolute % Difference |  |  |  |  |  |  |  | 7.4% |
| Range |  |  |  |  |  |  | -12.3% to -3.8% |  |

# Each Year the EPA Interlaboratory Performance Evaluation Program (PEP) Audit program selects 25% of our SLAMS and NAMS sites to use in their Performance Evaluation Program in order to establish a basis for determining bias of our network.  Once per quarter the PEP auditor installs a portable PM2.5 sampler within 4 feet of our site sampler at each of these selected sites, and they take a side-by-side 24 hour sample.  It is desirable to have the concentration values of our 24 hour sample within 10% of the independent PEP audit concentration value.  However, given the nature of the low concentration values we are recovering and the number of APCD/PEP concentration pairs that are not valid because they are below 6 g/m3, this is quite difficult.  The differences between our values and the independent PEP auditor values averaged over a year give us an indication of instrument bias for a specific site.   All sites averaged over the year gives us an indication of our network bias. Table 14 summarizes these bias results.

Table 14 - PM 2.5 Interlab Performance Evaluation Program Audit Results

| Site and ID | Quarter | Date | APCD Station | PEP Audit | Difference | BIAS | Temperature | Pressure |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Results (g/m3) | Results (g/m3) | in (g/m3) |  | (deg C) | (mmHg) |
| Longmont | 1 | 1/16/2001 | 9 | 8.07 | 0.93 | 11.5% | -6.1 | 639 |
| 080130003 | 2 | 5/22/2001 | 4.5 | 3.79 | 0.72 | N/A | 16.1 | 637 |
|  | 3 | 9/19/2001 | 7.6 | 6.20 | 1.40 | 22.6% | 16.8 | 638 |
|  | 4 | 10/31/2001 | 4.6 | 4.91 | -0.31 | N/A | 15.9 | 629 |
| Longmont Annual |  |  |  |  |  | 17.0% |  |  |
| Elbert | 1 | 1/10/2001 | 2.6 | 1.83 | 0.77 | N/A | 3.2 | 589 |
| 080390001 | 2 | 3/29/2001 | 4.4 | 4.95 | -0.55 | N/A | 1.2 | 584 |
|  | 3 | 5/16/2001 | 3.5 | 3.83 | -0.33 | N/A | 15.9 | 589 |
|  | 4 | 9/25/2001 | 5.6 | 5.20 | 0.40 | N/A | 17.1 | 594 |
| Elbert Annual |  |  |  |  |  |  |  |  |
| Co. Springs - Meadowland | 1 | 1/10/2001 | 6.8 | 6.20 | 0.60 | 9.7% | 4.5 | 598 |
| 080410008 | 2 | 5/16/2001 | 6 | 6.95 | -0.95 | -13.6% | 18 | 600 |
|  | 3 | 9/25/2001 | 7.7 | 8.11 | -0.41 | -5.1% | 28.3 | 605 |
|  | 4 | 12/6/2001 | 3.3 | 3.08 | 0.22 | N/A | 5.5 | 601 |
| Meadowland Annual |  |  |  |  |  | -3.0% |  |  |
| Greeley | 1 | 1/16/2001 | 9.8 | 7.41 | 2.39 | 32.3% | -6 | 635 |
| 081230008 | 2 | 5/22/2001 | 3.8 | 3.66 | 0.14 | N/A | 14.8 | 640 |
|  | 3 | 9/19/2001 | 8.4 | 6.91 | 1.49 | 21.6% | 29.7 | 638 |
|  | 4 | 10/31/2001 | 9.3 | 8.24 | 1.07 | 12.9% | 18.6 | 632 |
| Greeley Annual |  |  |  |  |  | 22.3% |  |  |
| CDPHE Annual BIAS |  |  |  |  |  | 12.10% |  |  |

3.4 Precision and Accuracy Program

The APCD is required to conduct a precision and accuracy assessment program for the particulate and lead monitoring networks.  As with the gaseous monitoring network, the accuracy audit program requires that at least 25% of the samplers are audited each quarter and that each monitor be audited at least once per year.

The precision testing program, for manual samplers, is different from that used for automated samplers.  The automated systems compare a known analyte concentration (actual value) with an analyzer response (indicated value) on a biweekly schedule.  There are no particulate concentration standards that can be used as "actual values," so the manual sampler precision program relies on comparison of collocated instruments.  Identical collocated instruments are operated at several sites (3 PM2.5, 2 TSP, 4 PM10, and 1 lead site), on an every sixth day sample schedule, and the resulting measurements are used as precision checks.  A sampler is selected as the primary sampler to represent the "actual values" and a collocated sampler provides the "indicated values."

The atmospheric concentration (g/m3) values determined in this precision test program are submitted to the EPA Aerometric Information Retrieval System (AIRS) database within 90 days of the end of each calendar quarter.  The results of accuracy audits are also submitted to AIRS at the end of each calendar quarter.  As in the precision test program, "actual" and "indicated" values, in this case sampler flow rates (m3/minute), are reported.  The precision and accuracy data submitted to AIRS are used to calculate precision probability limits, and accuracy probability limits.

The 2001 precision and accuracy probability limits for TSP and PM10 are presented in Table 15.  Precision confidence limits for PM2.5 are summarized in Table 16.  Note the series of calculations prescribed by the EPA in CFR Part 58 Appendix A revision 1997 for PM2.5 confidence limits is a different set of calculations than those used to establish probability limits for automated and other particulate methods.  Trends in APCD precision probability limits for PM10 and TSP pollutants over the past 11 years are presented graphically in Figures 7 and 8, and the past three years of PM2.5 data in Figure 9.  For PM10 and TSP the upper and lower 95% probability limits indicate the range of percent difference from the "actual value" that would include 95% of the "indicated values".  About 5% of the results of an individual precision or accuracy test would exceed these limits.  Ideally, the probability ranges would be very small, and centered around zero.  This indicates that the ambient data collected by the analyzer are both precise and accurate.   For PM2.5 the upper and lower 90% confidence interval limits represent the chi - squared distribution range where 90% of the calculated Coefficient of Variation values for individual precision tests would be found. 10% of the calculated Coefficient of Variation values would not fall within this range.  Ideally this range would be very small and aggregated around a very small number.

A significant increase in the limits of PM10 and TSP can be seen graphically for 2001 data in Figures 7 and 8.  This is not a reflection of our program.  This apparent increase in limits is due to a calculation change instituted in the 1997 CFR that we are now incorporating into our data assessment review.

No precision probability data was available for lead since none of the collocated values were at a high enough concentration to qualify as a valid precision pair.  Lead accuracy is a two-fold determination of TSP sampler flow rate accuracy and lab analytical accuracy.  The flow rates of TSP lead samplers are regularly audited as part of the particulate network operations.  The laboratory QA/QC program is discussed in section 4 of this 2001 QA Report.

Table 15 -2001 PM10 and TSP Precision and Accuracy Probability Limits

(95% Confidence Interval)

|  |  | Precision Probability Limits |  |  | Accuracy Probability Limits |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Quarter | # Valid Pairs | Lower PL | Upper PL | # of Audits | Lower PL | Upper PL |
| PM10 | 1 | 33 | -18% | 26% | 139 | -6% | 6% |
|  | 2 | 41 | -15% | 16% | 102 | -4% | 4% |
|  | 3 | 29 | -18% | 19% | 84 | -5% | 5% |
|  | 4 | 33 | -11% | 18% | 123 | -4% | 5% |
|  | Annual | 136 | -16% | 20% | 448 | -5% | 5% |
| TSP | 1 | 25 | -21% | 30% | 15 | -9% | 11% |
|  | 2 | 33 | -12% | 21% | 18 | -10% | 3% |
|  | 3 | 21 | -13% | 22% | 10 | -9% | 7% |
|  | 4 | 21 | -25% | 18% | 10 | -4% | 4% |
|  | Annual | 100 | -18% | 24% | 53 | -10% | 7% |

Table 16 – PM2.5 Precision Data Confidence Limits

| Site/Station Name | Quarter | Precision(CV) | Lower CL | Upper CL |
| --- | --- | --- | --- | --- |
| 2001 Precision |  |  |  |  |
| Adams City | 1 | 8.2 | 6.1 | 16.8 |
| Colorado Springs | 1 | 8.5 | 6.5 | 15.7 |
| Grand Junction | 1 | 6.7 | 5.0 | 14.4 |
| CDPHE | 1 | 7.9 | 6.6 | 11.3 |
| Adams City | 2 | 3.1 | 1.9 | 19.8 |
| Commerce City | 2 | 6.7 | 5.1 | 12.4 |
| Colorado Springs | 2 | 10.6 | 6.9 | 46.2 |
| Grand Junction | 2 | 4.1 | 2.7 | 14.2 |
| CDPHE | 2 | 6.8 | 5.5 | 10.3 |
| Commerce City | 3 | 4.3 | 3.1 | 9.7 |
| Colorado Springs | 3 | 4.8 | 3.6 | 9.8 |
| Grand Junction | 3 | 3.7 | 2.5 | 10.9 |
| CDPHE | 3 | 4.4 | 3.6 | 6.6 |
| Commerce City | 4 | 5.7 | 4.4 | 10.6 |
| CAMP | 4 | 3.2 | 2.5 | 5.4 |
| Grand Junction | 4 | 7.9 | 5.9 | 16.3 |
| CDPHE | 4 | 5.6 | 4.9 | 8.0 |
| Adams City | Annual | 7.4 | 5.7 | 13.7 |
| Commerce City | Annual | 5.8 | 5.0 | 8.5 |
| CAMP | Annual | 3.2 | 2.5 | 5.4 |
| Colorado Springs | Annual | 7.7 | 6.3 | 11.4 |
| Grand Junction | Annual | 6.4 | 5.4 | 9.7 |
| CDPHE | Annual | 6.3 | 5.7 | 7.1 |

Figure 7 - APCD Annual Comparison of PM10 Precision Probability Limits (1990-2001)

(95% Confidence Interval)

Figure 8 - APCD Annual Comparison of TSP Precision Probability Limits (1990-2001)

(95% Confidence Interval)

(95% Confidence Interval)

Figure 9 - APCD PM2.5 Precision Confidence Intervals (1999-2001)

(90% Confidence Interval)

3.5 Special Study Audits

The APCD often conducts short-term studies of particulate pollution in Colorado communities.  Special studies have been conducted during the past two winter seasons at Denver National Jewish Hospital.  These PM10 and PM2.5 samplers were audited several times over the duration of the study.  All audits conducted at this site passed EPA criteria.  Some short-term studies were conducted for EPA in Grand Junction.

3.6 Additional Audits

APCD will occasionally perform audits on particulate monitoring networks operated by other public and private entities in Colorado.  Although not part of the APCD’s monitoring network, the APCD has an interest in the quality of these data.  TSP audits were conducted at Cotter Corporation’s Canon City site.

3.7 Certification of Audit Standards

The TSP and PM10 audits and calibrations performed by APCD Technical Services staff are done using one of nine fixed or variable resistance orifices.  Until early 1992, the variable resistance orifices were certified at two-year intervals on a ROOTS meter having NIST (National Institute for Standards and Technology) traceability residing with EPA Region VIII offices in the Denver Federal Center.  EPA Region VIII closed their Air Quality Assurance Laboratory in 1992.  In early 1993, a ROOTS meter was donated to the APCD by a private monitoring company.  This meter is certified against a NIST traceable standard at least every three years.   APCD orifices are certified on an annual schedule using this meter.  In addition to these certifications, the APCD performs frequent certifications for private groups doing particulate monitoring in Colorado.

The PM2.5 audits and calibrations performed by APCD Technical Services staff are done with a flow transfer standard.  APCD should have a NIST traceable flow verification system in house to certify these instruments by the early part of 2002.  Currently we are sending these flow transfer standards to CEESI for NIST traceable verifications.

4.0 AIR POLLUTION LABORATORY QUALITY ASSURANCE

The Colorado Department of Public Health and Environment (CDPHE) Laboratory Division provides a number of analytical services to support the air monitoring work conducted by the Air Pollution Control Division.  These support services include gravimetric analysis of air filters used for particulate monitoring, and analysis of selected filters for several elemental and ionic species.  In 2001, analyses were conducted for lead (Pb), sulfate (SO4=), arsenic (As), and cadmium (Cd) in support of APCD monitoring programs.

The quality assurance program in the air pollution laboratory is comprised of a number of internal quality control checks and external audits.  Laboratory instrumentation undergoes routine checks and calibrations in accordance with the CDPHE Laboratory and Radiation Services (LARS) Quality Assurance Policy Manual and LARS Standard Operating Procedures (SOPs).

Internal quality control checks are performed for the gravimetric determination of particulate as well as the analytical determinations of metals and sulfates.  One important quality control check for gravimetric analysis of particulate filters involves having an independent analyst reweigh a randomly selected subset of 10% of each batch of filters.  If the indicated gross or tare weight varies from the previous weight by more than 0.2 mg on any sample the entire batch of filters is re-weighed.

Analytical quality control checks for lead and sulfate include duplicate analyses, spiked samples, spiked blanks, and the analysis of filter strips provided by the EPA as part of the National Performance Audit Program (NPAP).  These quality control checks are briefly summarized in Table 17.

Table 17 -Laboratory Quality Control Checks

| Quality Control Check | Procedure |
| --- | --- |
| Duplicates | A duplicate filter strip is cut from the exposed filter and run through the analysis with the 'official' filter strips. |
| Spikes | A strip cut from an exposed filter is analyzed, then treated with a known amount of pollutant and reanalyzed. |
| NPAP Interlabs | Filter strips provided by EPA are analyzed as unknowns for a single point assessment per batch of filters processed. |

4.1 Duplicate Samples

Laboratory precision estimates for lead and sulfate are derived from analysis of duplicate strips cut from the same TSP filter.  Duplicate strips are cut from every tenth filter and dissolved into an aqueous solution.  A percent difference is calculated as the absolute value of the difference (in micrograms per milliliter of filter extract) divided by the mean concentration of the two samples.  This percent difference is calculated for samples that contain the analyte above the practical quantitation level (PQL).  For lead, the PQL is 0.05 μg/ml of sample extract, and for sulfate the PQL is 0.5 μg/ml of sample extract.  The results of the duplicate analyses performed during 2001 are presented in Tables 19 and 20.  Analysis of these duplicate samples is an internal quality control check.  The results of these duplicate analyses are reported annually to the APCD, but are not submitted to the EPA AIRS database.

4.2 Spiked Samples

The CDPHE laboratory routinely performs analysis of spiked filter samples and filter blanks.  After the initial round of analysis, every tenth filter sample is spiked with a known amount of analyte and then reanalyzed.  The results of these spiked sample analyses are recorded on control charts.  Spiked sample analyte recoveries must fall within the laboratory control chart acceptance limits.  The results of these internal quality control checks are reported to the APCD annually.

In addition to the spiked sample analyses conducted on exposed filters, the CDPHE laboratory also prepares laboratory fortified blanks (or spiked blanks) by spiking clean filter strips with known amounts of analyte.  These audit sample strips are used to provide an assessment of analytical accuracy.  The laboratory currently analyzes filter strips at two lead concentrations: 40 μg/strip and 160 μg/strip.  The laboratory is required to audit at least three samples each calendar quarter at each of the two concentration ranges.  A summary of lead sample and blank spike results can be found in Table 18.

The analytical accuracy results for lead are reported by the APCD to the AIRS-AQS air quality database.  The analytical accuracy probability limits calculated from these lead analyses were presented in Table 15 of this report.

Although there is no federal requirement for analyzing sulfate strips, the laboratory also conducts monthly analyses of a 500 μg/strip sulfate fortified blank.  The results of the lead and sulfate spiked sample analyses performed by the CDPHE Laboratory and Radiation Services Division during 1999 and 2000 are presented in Table 20.  There is currently no provision to report the results of these sulfate analyses to the AIRS-AQS database.

4.3 EPA Interlaboratory Audits

The CDPHE Inorganic Laboratory participates in the EPA NPAP Interlaboratory audit program for lead.  In 1996, the EPA discontinued NPAP interlaboratory audits for sulfate.   Lead interlaboratory results are available for 3 quarters of 2001.  The EPA requests that laboratories with results exceeding 10% difference examine their measurement systems in order to reduce the size of the difference.  Laboratories with results exceeding 15% difference are asked to determine the cause and report any corrective actions to the NPAP Program Manager. Results for the lead interlaboratory audits are presented in Table 19.  The results of all the lead NPAP audits are within the 10% EPA criterion.

Table 18 - Laboratory Spike Recoveries for Lead

| 2001 | % Sample Spike Recovery |  |  |  |  | % Blank Spike Recovery |  | Average % Recovery |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| January | 91% | 98% | 90% |  |  | 95% | 102% | 95% |
| February | 95% | 91% | 99% | 96% |  | 95% | 99% | 96% |
| March | 97% | 96% | 86% | 93% |  | 91% | 94% | 93% |
| April | 102% | 101% | 103% |  |  | 91% | 96% | 99% |
| May | 92% | 94% | 107% | 109% |  | 98% | 100% | 100% |
| June | 117% | 98% |  |  |  | 98% | 100% | 103% |
| July | 121% | 99% |  |  |  | 97% | 99% | 104% |
| August | 95% | 101% |  |  |  | 97% | 99% | 98% |
| September | 107% | 97% | 112% |  |  | 100% | 100% | 103% |
| October | 93% | 92% | 96% |  |  | 99% | 99% | 96% |
| November | 94% | 90% |  |  |  | 99% | 99% | 96% |
| December | 89% | 104% | 98% | 95% |  | 104% | 103% | 99% |
| Summary |  |  |  |  |  |  |  | 98% |

Table 19 - Laboratory Duplicate Analyses and NPAP Results for Lead

|  | % Absolute Difference on Duplicate Analyses |  |  |  |  |  |  |  |  | NPAP Results |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 2001 |  |  |  |  |  |  |  |  |  |  |  |
| Q1 | 0.1% | 0.2% | 0.8% | 8.1% | 3.8% | 12.9% |  |  |  | -10.8% | -6.7% |
| Q2 | 8.7% | 6.5% | 3.9% | 2.1% |  |  |  |  |  | 1.4% | 0.6% |
| Q3 | 1.9% | 6.0% |  |  |  |  |  |  |  | 3.9% | 6.9% |
| Q4 | 5.8% | 0.3% | 9.5% | 1.7% | 12.7% | 9.1% | 10.4% | 9.9% | 0.6% | -13.5% | -12.8% |

Table 20 - Laboratory Quality Control for Sulfate

|  | Sample Spikes | Blank Spikes | Duplicates |
| --- | --- | --- | --- |
|  | % Recovery | % Recovery | % Difference |
| 2001 Results |  |  |  |
|  | 99% | 95% | 0% |
|  | 105% | 96% | 0% |
|  | 98% | 94% | 1% |
|  | 100% | 103% | 1% |
|  | 99% | 94% | 1% |
|  | 97% | 101% | 1% |
|  | 104% | 95% | 0% |
|  | 90% | 95% | 0% |
|  | 103% | 93% | 2% |
|  | 97% |  | 2% |
|  | 98% |  | 1% |
|  | 94% |  |  |
| 2001 Averages | 99% | 96% | 1% |

5.0 METEOROLOGICAL MONITORING QUALITY ASSURANCE

The Air Pollution Control Division operates a meteorological monitoring network which is used to provide information to support air quality modeling and air pollution episode forecasting.  These monitoring sites measure ambient temperature, wind speed, and wind direction at an elevation of 10 meters above the surrounding terrain.  During 2001 there were 19 permanent meteorological monitoring sites in the APCD network.

The APCD CMDSS unit is responsible for meteorological sensor calibrations and the data acquisition and telemetry systems at these sites.  The Quality Assurance Unit has been performing independent audits of the meteorological network since 1990.   Performance audit equipment and personnel are different than those used during calibration.  The results of these audits are reported to APCD management and staff, but there is currently no provision to submit these audit results to the AIRS-AQS database.

The audit procedures include an overall review of station operation, and an assessment of tower cross arm orientation using a surveyor's transit.   The wind direction sensor is audited by rotating the sensor to a known orientation.  Wind speed is audited by using a series of three synchronous motors to rotate the anemometer shaft.  The temperature sensor is audited by comparing the current station temperature with that measured using a NIST-traceable thermometer.  Data quality objectives for the APCD meteorological monitoring network include tower orientation within 5 C wind direction measurements within 5 checked at 9 positions, wind speed measurements within 5% of actual values checked at 5 speeds, and ambient temperature measurements within 2  C.

The APCD goal is to audit each site annually.  During 2001 thirteen accuracy audits were performed at the nineteen meteorological monitoring sites.  All parameters for these audits met APCD Data Quality Objectives with the following exceptions:

Temperature at Cripple Creek (3.2 C out)

Channel 11 (upper) temperature at Auraria (2.5 C out)

Temperature at Carriage (2.2 C out)

6.0 CORRECTIVE ACTION SUMMARY

This section summarizes all problems encountered within the APCD air pollutant monitoring programs over 2001 as well as all corrective actions that were taken to remedy these situations.

|  |  |  |  | QC/QA | % Full | Date of Last |
| --- | --- | --- | --- | --- | --- | --- |
| Parameter | Date | Site / Station Name | Corrective Action | Action | Scale | Valid QC, flow |
|  |  |  |  | Performed | Error | or Calibration |
| O3 | 1/8/2001 | NREL | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | variable | 1/8/2001 |
| O3 | 1/19/2001 | Welby | Maintenance / Calibration | Assessment | variable | N/A |
| O3 | 12/27/2001 | Welby | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | -14.9% | 12/21/2001 |
| O3 | 5/21/2001 | Welch | Maintenance / Calibration | Assessment | 11.9% | <15% |
| O3 | 10/9/2001 | Welch | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | -27.7% | 10/7/2001 |
| SO2 | 4/20/2001 | Welby | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | -11.0% | <15% |
| PM2.5 | 2/2/2001 | Commerce City | Immediate Calibration of Pressure/Recalculation of Data | Pressure | -12mm Hg | N/A |
| PM2.5 | 2/13/2001 | Durango | Seals Replaced | External Leak | Fail | N/A |
| PM2.5 | 3/30/2001 | Platteville | Immediate Calibration of Filter Temperature Probe | Temperature | 3.0 deg C | N/A |
| PM2.5 | 6/29/2001 | Parker | Seals replaced | Int/Ext Leaks | Fail | N/A |
| NOx | 1/25/2001 | CAMP | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | off scale | 1/11/2001 |
| NOx | 5/24/2001 | CAMP | Calibration / Data deleted back to last valid QC check point | Assessment | 19.9% | 5/21/2001 |
| NOx | 8/6/2001 | CAMP | Immediate Calibration | Assessment | -11.2% | <15% |
| NOx | 9/4/2001 | CAMP | Calibration | Accuracy Audit | 11.0% | <15% |
| NOx | 3/19/2001 | Rocky Flats SE | Calibration / Data deleted back to last valid QC check point | Accuracy Audit | -12.0% | <15% |
| NOx | 3/21/2001 | Rocky Flats SE | Calibration / Data deleted back to last valid QC check point | Assessment | -17.5% | 3/21/2001 |
| NOx | 6/28/2001 | Rocky Flats SE | Close out audit - instrument removed | Assessment | -13.4% | <15% |
| NOx | 3/15/2001 | Rocky Flats West | Maintenance / Calibration / Data deleted back to last valid QC check point | Accuracy Audit | 14.0% | 3/4/2001 |
| NOx | 3/16/2001 | Rocky Flats West | Immediate Calibration | Assessment | -14.7% | 3/4/2001 |
| NOx | 3/22/2001 | Rocky Flats West | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | -81.0% | 3/4/2001 |
| NOx | 4/19/2001 | Rocky Flats West | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | 13.7% | 3/4/2001 |
| NOx | 7/2/2001 | Rocky Flats West | Close out audit - instrument removed / data deleted back to last valid QC | Assessment | -16.8% | 6/30/2001 |
| NOx | 3/8/2001 | Welby | Calibration / Data deleted back to last valid QC check point | Assessment | -21.6% | 3/7/2001 |
| NOx | 6/22/2001 | Welby | Immediate Calibration | Assessment | -10.7% | <15% |
| NOx | 9/14/2001 | Welby | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | 70.0% | 9/10/2001 |
| NOx | 11/9/2001 | Welby | Maintenance / Calibration / Data deleted back to last valid QC check point | Assessment | 21.8% | 10/18/2001 |
| PM10 | 11/27/2001 | Alamosa #3 | Mass Flow Controller Adjusted | Design Flow | -17.0% | N/A |
| PM10 | 11/27/2001 | Alamosa #4 | Mass Flow Controller Adjsuted | Design Flow | 11.0% | N/A |
| PM10 | 12/13/2001 | Breckenridge | Mass Flow Controller Adjsuted | Design Flow | 11.0% | N/A |
| PM10 | 2/6/2001 | CO Springs - RBD (collocated) | Motor Replacement / Claibration | Accuracy Audit | 9.0% | <10% |
| PM10 | 2/2/2001 | Commerce City | Mass Flow Controller Adjsuted | Design Flow | 9.0% | <10% |
| PM10 | 3/12/2001 | Criple Creek #1 | Motor Replacement / Claibration | Accuracy Audit | 9.0% | <10% |
| PM10 | 6/19/2001 | Criplpe Creek #1 | Calibration | Accuracy Audit | -8.0% | <10% |
| PM10 | 3/12/2001 | Cripple Creek #4 | Motor Replacement / Claibration | Accuracy Audit | 8.0% | <10% |
| PM10 | 8/30/2001 | Denver Visitor Center | Mass Flow Controller Adjsuted | Design Flow | -10.0% | <10% |
| PM10 | 12/5/2001 | Denver Visitor Center | Calibration | Accuracy Audit | -9.0% | <10% |
| PM10 | 11/28/2001 | Durango Courthouse #1 | Mass Flow Controller Adjsuted | Design Flow | -11.0% | N/A |
| PM10 | 2/13/2001 | Durango Courthouse #2 | Incorrect Cassette Installation by Operator | Design Flow | -13.0% | N/A |
| PM10 | 3/14/2001 | Durango Courthouse #2 | Mass Flow Controller Adjsuted | Design Flow | 27.0% | N/A |
| PM10 | 2/13/2001 | Durango Platform #2 | Mass Flow Controller Adjsuted | Design Flow | 24.0% | N/A |
| PM10 | 3/14/2001 | Durango Platform #2 | Mass Flow Controller Adjsuted | Design Flow | 13.0% | N/A |
| PM10 | 5/22/2001 | Durango Platform #3 | Mass Flow Controller Adjsuted | Design Flow | 11.0% | N/A |
| PM10 | 3/14/2001 | Durango School #1 | Mass Flow Controller Adjsuted | Design Flow | -11.0% | N/A |
| PM10 | 5/22/2001 | Durango School #1 | Incorrect Cassette Installation by Operator | Design Flow | -29.0% | N/A |
| PM10 | 3/14/2001 | Durango School #2 | Mass Flow Controller Adjsuted | Design Flow | -14.0% | N/A |
| PM10 | 8/7/2001 | Ft. Collins - Courthouse | Maintenance / Calibration/ Data deleted back to last valid Calibration | Accuracy Audit | 100.0% | 3/30/2001 |
| PM10 | 3/6/2001 | Ft. Collins - School | Calibration | Accuracy Audit | -9.0% | <10% |
| PM10 | 6/22/2001 | Hygiene #2 | Calibration | Accuracy Audit | -8.0% | <10% |
| PM10 | 3/6/2001 | Hygiene #3 | Calibration | Accuracy Audit | 8.0% | <10% |
| PM10 | 6/18/2001 | Lamar Municipal Blgd #1 | Mass Flow Controller Adjsuted | Design Flow | 14.0% | N/A |
| PM10 | 10/15/2001 | Lamar Municipal Blgd #1 | Mass Flow Controller Adjsuted | Design Flow | -12.0% | N/A |
| PM10 | 6/18/2001 | Lamar Power Plant #3 | Mass Flow Controller Adjsuted | Design Flow | 11.0% | N/A |
| PM10 | 5/21/2001 | Pagosa #1 | Mass Flow Controller Adjsuted | Design Flow | 13.0% | N/A |
| PM10 | 5/21/2001 | Pagosa #3 | Mass Flow Controller Adjsuted | Design Flow | 12.0% | N/A |
| TSP | 8/3/2001 | CAMP | Outside Warning Ranges | Design Flow | -11.0% | N/A |
| TSP | 10/1/2001 | CAMP | Outside Warning Ranges | Design Flow | -11.0% | N/A |
| TSP | 8/15/2001 | Grand Junction Traffic (collocated) | Outside Warning Ranges | Design Flow | 11.0% | N/A |
| TSP | 11/12/2001 | Grand Junction Traffic (primary) | Outside Warning Ranges | Design Flow | -14.0% | N/A |
| TSP | 4/18/2001 | Commerce City | Motor Replaced/Calibrated/Data deleted back to last valid Calibration | Accuracy Audit | -13.0% | 1/9/2001 |

APPENDIX A - 2001 Site List

2001 Site List

| SITE / STATION NAME | ADDRESS | COUNTY | AIRS CODE | O3 | CO | SO2 | NOX | PM10 | TSP | Pb | PM2.5 | Met | Other |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Adams City (primary) | 4301 E. 72nd Ave. | Adams | 08-001-0001 |  |  |  |  | 1/1 | P | S | S |  |  |
| Adams City (colocated) | 4301 E. 72nd Ave. | Adams | 08-001-0001 |  |  |  |  |  |  |  | 1/6 |  |  |
| Globeville Clinicare | 5400 Washington St. | Adams | 08-001-0005 |  |  |  |  |  | 1/3-1/6 | P |  |  |  |
| Commerce City (primary) | 7101 Birch St. | Adams | 08-001-0006 |  |  |  |  | 1/1 | P | S | S |  | x |
| Commerce City (colocated) | 7101 Birch St. | Adams | 08-001-0006 |  |  |  |  |  |  |  | 1/6 |  |  |
| South Adams Pump Station | 5580 Niagra St. | Adams | 08-001-0007 |  |  | S |  |  |  |  |  |  |  |
| Brighton | 22 S. 4th Ave. | Adams | 08-001-2002 |  |  |  |  | 1/3-1/6 |  |  |  |  |  |
| Welby | 78th Ave. & Steele St. | Adams | 08-001-3001 | S | S | N | N | 1/6 |  |  |  | x | x |
| Alamosa | 369 Richardson Ave. | Alamosa | 08-003-0001 |  |  |  |  | 1/1 |  |  |  |  |  |
| Highlands | 8100 S. University | Arapahoe | 08-005-0002 | S |  |  |  |  |  |  |  | x |  |
| Arapahoe Community College | 6190 S. Santa Fe Dr. | Arapahoe | 08-005-0005 |  |  |  |  |  |  |  | S |  |  |
| Pagosa Springs | Middle School | Archuleta | 08-007-0001 |  |  |  |  | 1/1 |  |  | P |  |  |
| Pagosa Springs | 486 San Juan Ave. | Archuleta | 08-007-0002 |  |  |  |  | 1/1 |  |  | P |  |  |
| Longmont (particulate) | 3rd Ave. & Kimbark St. | Boulder | 08-013-0003 |  |  |  |  | 1/3 |  |  | S |  |  |
| Longmont (gaseous) | 440 Main St. | Boulder | 08-013-0009 |  | S |  |  |  |  |  |  |  |  |
| Boulder - YMCA | 2150 28th St. | Boulder | 08-013-0010 |  | S |  |  |  |  |  |  |  |  |
| Boulder - S. Boulder Creek | 14051/2 S. Foothills Hwy. | Boulder | 08-013-0011 | S |  |  |  |  |  |  |  |  |  |
| Boulder Chamber of Commerce | 2440 Pearl St. | Boulder | 08-013-0012 |  |  |  |  | 1/6 |  |  | S |  |  |
| Hygiene | 7024 Colorado Hwy 66 | Boulder | 08-013-0013 |  |  |  |  | 1/1 |  |  |  |  |  |
| Delta | 560 Dodge St. | Delta | 08-029-0004 |  |  |  |  | 1/6 |  |  |  |  |  |
| Paonia | Middle School | Delta | 08-029-0005 |  |  |  |  | 1/6 |  |  |  |  |  |
| Hotchkiss | 222 W. Bridge St. | Delta | 08-029-0006 |  |  |  |  | 1/6 |  |  |  |  |  |
| Denver - CAMP (primary) | 2105 Broadway | Denver | 08-031-0002 |  | S | N | N | 1/6 | P | N | S1/1 | x | x |
| Denver - CAMP (collocated) | 2105 Broadway | Denver | 08-031-0002 |  |  |  |  | 1/6 |  |  | 1/6 |  |  |
| Denver - NJH | 14th Ave. & Albion St. | Denver | 08-031-0013 |  | S |  |  |  |  |  |  |  |  |
| Denver - Carriage | 23rd Ave. & Julian St. | Denver | 08-031-0014 | S | S |  |  |  |  |  |  | x |  |
| Denver Gates (primary) | 1050 S. Broadway | Denver | 08-031-0015 |  |  |  |  | 1/6 | P | N |  |  | x |
| Denver Gates (collocated) | 1050 S. Broadway | Denver | 08-031-0015 |  |  |  |  | 1/6 | P | N |  |  |  |
| Denver Visitor Center | 225 W Colfax Ave. | Denver | 08-031-0017 |  |  |  |  | 1/1 |  |  |  |  |  |
| Denver Firehouse #6 | 1300 Blake St. | Denver | 08-031-0019 |  | S |  |  |  |  |  |  |  |  |
| Auraria Met | Auraria Campus Parking Lot | Denver | 08-031-0021 |  |  |  |  |  |  |  |  | x |  |
| Denver - LARS | 8100 Lorwy Blvd. | Denver | 08-031-0022 |  |  |  |  | 1/3 |  |  |  |  |  |
| Castle Rock | 310 3rd St. | Douglas | 08-035-0001 |  |  |  |  | 1/6 |  |  |  |  |  |
| Chatfield Reservoir | Roxborough Park Rd. | Douglas | 08-035-0002 | S |  |  |  |  |  |  |  |  |  |
| Parker | Library | Douglas | 08-035-0003 |  |  |  |  |  |  |  | P |  |  |
| Vail - WSD | 846 Forest Rd. | Eagle | 08-037-0007 |  |  |  |  | 1/6 |  |  |  |  |  |
| Elbert | Ben Kelly Rd. & Colo 5 & 98 | Elbert | 08-039-0001 |  |  |  |  |  |  |  | S |  |  |
| Colorado Springs - GLEN | I-25 & Uintah Ave. | El Paso | 08-041-0006 |  | S |  |  |  |  |  |  |  |  |
| Colorado Springs - Meadowlands | 3730 Meadowlands Blvd. | El Paso | 08-041-0008 |  |  |  |  | 1/3 |  |  | S |  |  |
| Colorado Springs-RBD (primary) | 101 Costillia | El Paso | 08-041-0011 |  |  |  |  | 1/6 | P | S | S |  |  |
| Colorado Springs-RBD(collocated) | 101 Costillia | El Paso | 08-041-0011 |  |  |  |  | 1/6 |  |  | 1/6 |  |  |
| C.S. - US Air Force Academy | Road 640, USAF Academy | El Paso | 08-041-0013 | N |  |  |  |  |  |  |  |  |  |
| Colorado Springs - Highway 24 | 69 W. U.S. Hwy 24 | El Paso | 08-041-0015 |  | S |  |  |  |  |  |  |  |  |
| Canon City | 7th Ave. & Macon St. | Freemont | 08-043-0001 |  |  |  |  | 1/6 |  |  |  |  |  |
| Parachute | 100 E. 2nd Ave. | Garfield | 08-045-0005 |  |  |  |  | 1/3 |  |  |  |  |  |
| Rifle | 200 W. 3rd St. | Garfield | 08-045-0006 |  |  |  |  | 1/6 |  |  |  |  |  |
| Glenwood Springs | 806 Cooper Ave | Garfield | 08-045-1002 |  |  |  |  | 1/6 |  |  |  |  |  |
| Crested Butte | Colo. Hwy 135 & Whiterock | Gunnison | 08-051-0004 |  |  |  |  | 1/3 |  |  |  |  |  |
| Mt. Crested Butte | Town Center | Gunnison | 08-051-0005 |  |  |  |  | 1/1 |  |  | P |  |  |
| Gunnison | 221 N. Wisconsin | Gunnison | 08-051-0006 |  |  |  |  | 1/3 |  |  |  |  |  |
| Arvada | 57th Ave & Garrison St. | Jefferson | 08-059-0002 | N | S |  |  |  |  |  |  | x |  |
| Welch | 12400 W. U.S. Hwy 285 | Jefferson | 08-059-0005 | S |  |  |  |  |  |  |  | x |  |
| Rocky Flats X-1 (N) (primary) | 16600 Colo 128 | Jefferson | 08-059-0006 | P |  |  |  | 1/6 | P |  |  | x |  |
| Rocky Flats X-1 (N) (collocated) | 16600 Colo 128 | Jefferson | 08-059-0006 |  |  |  |  |  | P |  |  |  |  |
| Rocky Flats X-2 (NE) (primary) | 11501 Indiana St. | Jefferson | 08-059-0007 |  |  |  |  | 1/6 | P |  |  | x |  |
| Rocky Flats X-2 (NE) (collocated) | 11501 Indiana St. | Jefferson | 08-059-0007 |  |  |  |  | 1/6 |  |  |  |  |  |
| Rocky Flats X-3 (SE) | 9901 Indiana St. | Jefferson | 08-059-0008 |  |  |  | P | 1/6 | P |  |  | x |  |
| Rocky Flats X-4 (S) | 18000 W. Hwy 72 | Jefferson | 08-059-0009 |  |  |  |  | 1/6 | P |  |  | x |  |
| Rocky Flats X-5 (W) | 11190 N Hwy 93 | Jefferson | 08-059-0010 |  |  |  | P | 1/6 | P |  |  | x |  |
| NREL | 20th Ave. & Quaker St. | Jefferson | 08-059-0011 | S |  |  |  |  |  |  |  |  |  |
| Leadville | 510 Harrison St. | Lake | 08-065-0001 |  |  |  |  |  | P | S |  |  |  |
| Durango - Platform | 277 3rd Ave. | La Platta | 08-067-0007 |  |  |  |  | 1/1 |  |  |  |  |  |
| Durango - Park School | 623 E. 5th Ave. | La Platta | 08-067-0008 |  |  |  |  | 1/2 |  |  | P |  | x |
| Durango - Courthouse | 1060 E. 2nd Ave. | La Platta | 08-067-1001 |  |  |  |  | 1/2 |  |  |  |  |  |
| Ft Collins - Courthouse | 200 W. Oak St. | Larimer | 08-069-0001 |  |  |  |  | 1/3 |  |  |  |  |  |
| Ft. Collins - CSU | 251 Edison Dr. | Larimer | 08-069-0009 |  |  |  |  | 1/3 |  |  | S |  |  |
| Ft. Collins (gaseous) | 708 S. Mason St. | Larimer | 08-069-1004 | S | S |  |  |  |  |  |  | x |  |
| Grand Junct - Health Dpt.(primary) | 515 Paterson Rd. | Mesa | 08-077-0003 |  |  |  |  | 1/6 | P |  | S |  |  |
| Grand Junct - Health Dpt.(colloc) | 515 Paterson Rd. | Mesa | 08-077-0003 |  |  |  |  |  |  |  | 1/6 |  |  |
| Grand Junction - Stocker | 12th St. & North Ave. | Mesa | 08-077-0014 |  | S |  |  | 1/6 |  |  |  | x | x |
| Grand Junction - Traffic | 924 4th Ave. | Mesa | 08-077-0016 |  |  |  |  |  | P |  |  |  |  |
| Montrose | 125 S. Townsand Ave. | Montrose | 08-085-0003 |  |  |  |  | 1/6 |  |  |  |  |  |
| Olathe | 327 4th St. | Montrose | 08-085-0004 |  |  |  |  | 1/3 |  |  |  |  |  |
| Aspen | 420 Main St. | Pitkin | 08-097-0004 |  |  |  |  | 1/3 |  |  |  |  | x |
| Lamar Power Plant | 100 2nd Ave. | Prowers | 08-099-0001 |  |  |  |  | 1/1 |  |  |  |  |  |
| Lamar Municipal Bldg | 104 E. Parmenter St. | Prowers | 08-099-0002 |  |  |  |  | 1/1 |  |  |  | x |  |
| Pueblo Public Works | 211 D St. | Pueblo | 08-101-0012 |  |  |  |  | 1/3 |  |  | S |  |  |
| Steamboat Springs | 136 6th St. | Routt | 08-107-0003 |  |  |  |  | 1/1 |  |  | P |  |  |
| Steamboat Springs Met | 137 10th St. | Routt | 08-107-0008 |  |  |  |  |  |  |  |  | x |  |
| Telluride | 333 W. Colorado Ave. | San Miguel | 08-113-0004 |  |  |  |  | 1/3 |  |  | P |  | x |
| Telluride Met | Coonskin Parking Lot | San Miguel | 08-113-0005 |  |  |  |  |  |  |  |  | x |  |
| Breckenridge | County Justice Cntr | Summit | 08-117-0002 |  |  |  |  | 1/3-1/1 |  |  |  |  |  |
| Silverthorne | 151 4th St. | Summit | 08-117-0003 |  |  |  |  | 1/6 |  |  |  |  |  |
| Silverthorne Rec Center | 430 Rainbow Dr. | Summit | 08-117-0004 |  |  |  |  | 1/6 |  |  |  |  |  |
| Cripple Creek | Bennet Ave. & 2nd St. | Teller | 08-119-0001 |  |  |  |  | 1/1 |  |  |  |  |  |
| Cripple Creek Met | S. 2nd St. & Warren Ave. | Teller | 08-119-0002 |  |  |  |  |  |  |  |  | x |  |
| Greeley (particulate) | 1516 Hospital Rd. | Weld | 08-123-0006 |  |  |  |  | 1/3 |  |  | S |  |  |
| Greely (gaseous) | 811 15th St. | Weld | 08-123-0007 | S | S |  |  |  |  |  |  |  |  |
| Platteville | Valley Middle School | Weld | 08-123-0008 |  |  |  |  |  |  |  | S |  |  |
| NOTES: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| S = SLAMS site |  |  |  |  |  |  |  |  |  |  |  |  |  |
| N = NAMS site |  |  |  |  |  |  |  |  |  |  |  |  |  |
| P = Special Purpose Monitoring site |  |  |  |  |  |  |  |  |  |  |  |  |  |
| 1/1, 1/2, 1/3, 1/6 refers to sampling frequency |  |  |  |  |  |  |  |  |  |  |  |  |  |
| All TSP Sites are 1/6 with the exception of Globeville |  |  |  |  |  |  |  |  |  |  |  |  |  |
| All PM2.5 SPM sites are 1/6 and all PM2.5 SLAMS sites are 1/3 with the exception of CAMP |  |  |  |  |  |  |  |  |  |  |  |  |  |
| OTHER samplers include TEOMs, Beta Gauges, and Speciation Samplers |  |  |  |  |  |  |  |  |  |  |  |  |  |

For more detailed information about specific monitoring sites, please refer to the APCD Annual Network Review.